Landmine and UXO Victims in Lebanon: The Future Isn’t as Grim Anymore

A Feasibility Study on the Best Solution for Upper and Lower Limb Prosthetics

To: Dr. Zane Sinno

By: Karim El Khatib, Nour El Khatib and Freddy Vartabedian

Course: English 206

Date: December 23, 2012

Abstract

Prosthetic limbs are a basic need for landmine and UXO surviving victims in Lebanon. Current prosthetic limbs are divided into conventional or unpowered artificial limbs and advanced and/or powered bionic limbs. The purpose of this report is to determine the extent to which highly advanced and high-cost prosthetic limbs can change the lives of Lebanese landmine and UXO victims and the feasibility of these advanced limbs in Lebanon.

To achieve our purpose, we have analyzed, in detail, two lines of prosthesis solutions, a conventional and an advanced one, for each of below knee, above knee and upper extremity cases. Our evaluation of the chosen prosthetic limbs is based on a detailed review of the literature on the landmine and landmine victims’ situation in Lebanon, a detailed review of the literature on the different types of chosen prosthetic limbs for this report, as well as a field research including the analysis of media and users comments and interviews.

Our evaluation is based on five carefully chosen criteria for each of lower limb prosthetics and upper extremity ones. The report includes an assessment of below knee prosthetics, mainly the low cost and conventional SACH foot and Jaipur foot, as well as the high-cost and highly advanced BiOM ankle. It also includes an evaluation of above knee prosthetics, including the simple and low cost single axis constant friction knee unit, the recent and low cost polycentric Jaipur knee and the highly advanced bionic leg, the BiOM AK. The report also addresses upper extremity prosthetics and includes an evaluation of the conventional body-powered arm, the highly advanced Luke arm coupled with the revolutionary Targeted Muscle Reinnervation surgical technique, and the recent and brain-controlled Artificial Muscle Operated arm.

Our results reveal the various advantages and disadvantages associated with the use of each of the mentioned types, with the superiority of advanced prosthetics in most aspects. Our discussion of the obtained results highlights the main advantages and disadvantages with a focus on the extent to which each type of prosthetic limb can enhance, change or even restore normalized functionality for amputees. Our conclusions reveal the significant capacity of the advanced prosthetics in restoring productivity and enhancing or even normalizing the amputee’s quality of life, thus leading us to reach a recommendation towards the use of advanced prosthetic limbs to restore a normal life for landmine and UXO victims in Lebanon

Table of Contents

1. Introduction1

1.1. Landmine and Unexploded Ordnance Victims in Lebanon and their Need for Prosthetic Limbs 1

1.2. Description of Landmine and Unexploded Ordnance victims’ status in Lebanon 1

1.2.1. The effects of landmines and unexploded ordnance on the socio-economic level1

1.2.2. The effects of landmines and unexploded ordnance on the human level2

1.2.3. The need to restore productivity through prosthetic limbs2

1.3. Possible Solutions 2

1.3.1. Conventional and low cost prosthetic limbs 2

1.3.1.1. Conventional prosthetic feet for below knee amputees2

1.3.1.2. Conventional prosthetic legs for above knee amputees 3

1.3.1.2. Conventional upper extremity prosthetics3

1.3.2. Innovative smart or intuitively controlled prosthetics3

1.3.2.1. The BiOM ankle for below knee amputees3

1.3.2.2. The BiOM AK for above knee amputees3

1.3.2.3. Myoelectric arms coupled with Targeted Muscle Reinnervation3

1.3.2.4. Artificial Muscle Operated arm4

1.4. Purpose and Overview: Study of the feasibility of the two possible solution 4

2. Method4

2.1. Assessment Criteria 4

2.1.1. Assessment criteria for prosthetic feet and legs4

1. Versatility of mobility and associated stability 5

2. Walking speeds and accompanying gait quality5

3. Metabolic demands.

4.  Practicality of usage5

5. Cost5

2.1.2. Assessment criteria for prosthetic arms5

1. Elbow, wrist and fingers: range of motion and associated scope of functionality 5

2. Control and powering system 5

3.Weight and Comfort Factors 5

4. Practicality of usage…………………………………………………………………………………………………………….    5

5. Cost 6

2.2. Procedure 6

2.2.1. Review of literature6

2.2.1.1. Review of literature on landmines 6

2.2.1.2. Review of literature on prosthetics 6

2.2.1.3. Final Review6

2.2.2. Field research6

2.2.2.1. Analysis of User Comments 6

2.2.2.2. Webinar 6

2.2.2.3. Interview with Dr. N. Sabah, professor of biomedical engineering at AUB 7

3. Results7

3.1. An overview of the human gait cycle 7

3.2. Below knee prosthetics10

3.2.1. General overview of prosthetic feet10

3.2.2. Conventional prosthetic feet10

3.2.2.1. The SACH foot 11

3.2.2.1.1. Versatility of mobility and associated stability 11

3.2.2.1.2. Walking speeds and accompanying gait quality  12

3.2.2.1.3. Metabolic demands 13

3.2.2.1.4. Practicality of usage 13

3.2.2.1.5. Cost 14

3.2.2.2. The Jaipur foot 14

3.2.2.2.1. Versatility of mobility and associated stability 14

3.2.2.2.2. Walking speeds and accompanying gait quality 15

3.2.2.2.3. Metabolic demands 16

3.2.2.2.4. Practicality of usage 17

3.2.2.2.5. Cost 17

3.2.3. The Power Foot: BiOM ankle18

3.2.3.1. Versatility of mobility and associated stability 19

3.2.3.2. Walking speeds and accompanying gait quality 20

3.2.3.3. Metabolic demands 21

3.2.3.4. Practicality of usage 22

3.2.3.5. Cost 23

3.3. Above knee prosthetics23

3.3.1. AK prosthetic components, parts and functioning modes23

3.3.2. Low cost conventional or mechanical AK prosthetics26

3.3.2.1. Single axis constant friction knee 26

3.2.2.1.1. Versatility of mobility and associated stability 27

3.2.2.1.2. Walking speeds and accompanying gait quality  27

3.2.2.1.3. Metabolic demands 29

3.2.2.1.4. Practicality of usage 29

3.2.2.1.5. Cost 30

3.3.2.2. The Jaipur Knee 30

3.3.2.2.1. Versatility of mobility and associated stability 30

3.3.2.2.2. Walking speeds and accompanying gait quality  32

3.3.2.2.3. Metabolic demands 33

3.3.2.2.4. Practicality of usage 34

3.3.2.2.5. Cost 34

3.3.3. The BiOM AK: The X2 Bionic Knee4434

3.3.3.1. Versatility of mobility and associated stability 35

3.3.3.2. Walking speeds and accompanying gait quality 37

3.3.3.3. Metabolic demands 38

3.3.3.4. Practicality of usage 39

3.3.3.5. Cost 39

3.4. Upper extremity prosthetics39

3.4.1. Conventional or body-powered prosthetic arms 40

3.4.1.1. Elbow, wrist and fingers: range of motion and associated scope of functionality 43

3.4.1.2. Control and powering factors 44

3.4.1.3. Weight and comfort factors 45

3.4.1.4. Practicality of usage 45

3.4.1.5. Cost 46

3.4.2. The myoelectric arm coupled with Targeted Muscle Reinnervation (TMR): the Luke arm46

3.4.2.1. The myoelectric arm 46

3.4.2.2. Targeted Muscle Reinnervation 47

3.4.2.3. The Luke arm 48

3.4.2.3.1. Elbow, wrist and fingers: range of motion and associated scope of functionality 49

3.4.2.3.2. Control and powering system 50

3.4.2.3.3. Weight and Comfort Factors 51

3.4.2.3.4. Practicality of usage 51

3.4.2.3.5. Cost 52

3.4.3. The Artificial Muscle Operated Arm (AMO)52

3.4.3.1. Elbow, wrist and fingers: range of motion and associated scope of functionality 53

3.4.3.2. Control and powering system 54

3.4.3.3. Weight and comfort factors 55

3.4.3.4. Practicality of usage 55

3.4.3.5. Cost 55

4. Discussion56

4.1. Below knee prosthetic feet 56

4.1.1. The SACH Foot56

4.1.1.1. Advantages 56

4.1.1.2. Disadvantages 57

4.1.1.3. Conclusion 58

4.1.2. The Jaipur Foot58

4.1.2.1. Advantages 58

4.1.2.2. Disadvantages 60

4.1.2.3. Conclusion 60

4.1.3. The Power Foot: BiOM Ankle61

4.1.3.1. Advantages 61

4.1.3.2. Disadvantages 62

4.1.3.3. Conclusion 63

4.2. Above knee prosthetic legs63

4.2.1. The Single Axis Constant Friction Knee63

4.2.1.1. Advantages 64

4.2.1.2. Disadvantages 64

4.2.1.3. Conclusion 65

4.2.2. The Jaipur Knee65

4.2.2.1. Advantages 66

4.2.2.2. Disadvantages 67

4.2.2.3. Conclusion 67

4.2.3. The BiOM AK: The X2 bionic knee 68

4.2.3.1. Advantages 68

4.2.3.2. Disadvantages 70

4.2.3.3. Conclusion …………………………………………………………………………………………………………………..70

4.3. Prosthetic Arms70

4.3.1. Body-Powered Arms71

4.3.1.1. Advantages 71

4.3.1.2. Disadvantages 72

4.3.1.3. Conclusion 72

4.3.2. The Luke Arm73

4.3.2.1. Advantages 73

4.3.2.2. Disadvantages 74

4.3.2.3. Conclusion 75

4.3.3. The Artificial Muscle Operated Arm 75

4.3.3.1. Advantages 75

4.3.3.2. Disadvantages 76

4.3.3.3. Conclusion 77

5. Conclusions and Recommendations77

6. References…………………………………………………………………………………………………………………………………………79

List of Tables, Graphs, and Figures

I-Tables

Table 1: Stages of the human gait cycle8

Table 2: Basic Foot movements9

Table 3: Articulations of the Jaipur foot 15

Table 4: Comparison of prosthetic and normal gait features 28

Table 5: TMR50

II-Graphs

Graph 1: Preferred Velocity20

Graph 2: Metabolic Cost of Transport 21

III-Figures

Figure 1: Landmine1

Figure 2: UXO1

Figure 3: The phases of the human gait cycle 8

Figure 4: The SACH Foot11

Figure 5: The SACH Foot 11

Figure 6: A sagittal section of the Jaipur foot 14

Figure 7: The Jaipur leg 15

Figure 8: The Power Foot BiOM 18

Figure 9: The structure of BiOM 19

Figure 10: The android application for tuning the BiOM 21

Figure 11: The motion of the lower leg in swing phase 24

Figure 12: The motion of the lower leg in swing phase 25

Figure 13: A single axis knee unit 26

Figure 14: The rotational trajectory of the center of the polycentric knee 27

Figure 15: The Jaipur knee30

Figure 16: The rotational trajectory of the center of the polycentric knee 31

Figure 17: The Jaipur Gait33

Figure 18: The BiOM AK35

Figure 19: Walter Reed X2 Knee Test 36

Figure 20: Traditional functional upper limb prosthesis with cable-controlled split hooks for hands 40

Figure 21: Terminal devices41

Figure 22: Below elbow prosthetic arm 41

Figure 23: Above elbow prosthetic 42

Figure 24: Bilateral shoulder disarticulation prosthesis and special harness 43

Figure 25: TMR47

Figure 26: The Luke Arm48

Figure 27: The Luke Arm49

Figure 28: AMO: Headset and arm53

Figure 29: The AMO arm54

- Introduction

  - Landmine and Unexploded Ordnance Victims in Lebanon and their Need for Prosthetic Limbs

This report addresses the situation of a marginalized fraction of the Lebanese population, namely the surviving victims of landmines and unexploded ordnance, and their need for efficient prosthetic limbs that can go beyond the level of mere limb replacement to achieve the restoration of a normal life for these victims.

  - Description of the Landmine and Unexploded Ordnance victims’ status in Lebanon

In Lebanon, landmines and unexploded ordnances have and still pose a major problem for a large fraction of the population. According to ICBL and as of 2005,

Figure 1: Landmine ©thoughtleadershipleverage.com

despite the huge demining accomplishments, around 550000                                                  landmines lie implanted in around 116.4 Km2 of Lebanese soil, waiting in silence to injure and maim without mercy man, woman and child alike (Yammine, 2007). Cluster bombs, on the other hand, which are a legacy of the last three days of the July 2006 war, were approximated in 2007 to be around one million in number infesting around 26% of agricultural land in South Lebanon including land that was previously cleared of landmines (Yammine, 2007).

Though clear figures of the exact number of landmines and unexploded ordnance (UXO) today are hard to find, their devastating effects linger on new and old victims, their families, their land, their socio-economic status and consequently on the economy of Lebanon in general. According to an LMRC report, the total number of civilian landmine and cluster bomb victims in Lebanon until June 2008 is 3732 including 2732 survivors (Balamand, 2008).

Figure 2:UXO ©electronic intifada.net

1.2.1. The effects of Landmines and Unexploded Ordnance on the Socio-economic Level

The survivors face the extreme difficulty of social integration due to the loss of a leg, both legs and sometimes even the arms, leading to loss of mobility and as a result, to a cycle of dependence and eventual poverty which forces them to the margins of life in all dimensions. In one study, where the injured were assumed to retain 50% of their productivity over the period of 20 years following the accident, the total loss per amputee was estimated to be $58530$, while the loss per each death was estimated to be $107060$ (Yammine, 2007 ). These figures, along with the fact that the people mostly affected by landmines and cluster bombs work in farming and the presence of these silent killers in their lands make it impossible for them to make full use of their resources and stand in the way for new investments (Yammine, 2007), allow us to conclude that landmines and cluster bombs result in devastating economic loss for the victims, their families and their communities even long after the war is over.

1.2.2. The Effects of Landmines and Unexploded Ordnance on the Human Level

On the human level, the surviving victims suffer a long chain of events following the trauma that results in a vicious cycle of loss of limbs essential for activity and productivity, poverty and loss of self-esteem, to the degree of isolation. Some even end up as beggars, while others are stigmatized by the community and most are eventually considered a burden and neglected. In the case of women who in most of these communities are also farmers or help their husbands in various ways, the effect is no different and some are even deserted by their husbands, while in the case of family providers such as father or brother, the whole family suffers extreme poverty and social marginalization which could stretch even for a lifetime. The extreme poverty also stands as an impediment between these victims and the seeking of professional medical assistance.

1.2.3. The Need to Restore Productivity through Prosthetic Limbs

Evidently, landmine and UXO victims need a way to restore, to a reasonable degree, their normal lives and to reintegrate in society and resume their lives as full members of the community. Restoring productivity is the key to break the vicious cycle of their lives and that is only possible by giving them mobility or lost limb activity, or in other words, replacing the lost limbs with artificial ones. Prosthetics is the answer to these victims to regain control over their lives. Article six of the Antipersonnel Mine Ban Treaty of ICBL (Ottawa Convention) states that, “Each State Party in a position to do so shall provide assistance for the care and rehabilitation, and social and economic reintegration of mine victims and for mine awareness programs” (E-Mine, 2009, p.1). A similar one addresses cluster munitions victims, their families and communities. Also, the 2003 protocol on Explosive Remnants of War calls for assistance for victims (ICRC, 2005).

Given all the preceding factors, it is justified to conclude that prosthetics are a basic need in such communities and they are the primary step in the process of rehabilitation and socio-economic reintegration of landmine and UXO victims, as well as the basic step for their psychological well-being and the restoration of a dignified life for them and their families.

  - Solutions:

In our project we proposed and examined a twofold solution comprising the use of conventional and low cost prosthetic limbs and the highly sophisticated recent technology of intuitively or thought controlled prosthetics using Targeted Muscle Reinnervation and the Artificial Muscle Operated arm for upper extremities and the intelligent leg, a combination of the BiOM ankle and the Otto Bock X2 knee, for leg amputees.

    - Conventional and Low Cost Prosthetic Limbs

In developing countries like Lebanon, where the cost of a prosthetic is a major issue and victims of landmines and UXOs lack help from the government and insurance coverage of prosthetics is inexistent or only partial, durable, efficient and low cost prosthetics present a major solution for amputees. In answer to this need we examined the following solutions:

1.3.1.1. Conventional Prosthetic Feet for Below Knee Amputees

We reviewed two highly recommended conventional prosthetic feet, the SACH foot and the Jaipur foot for below knee amputees and we will attempt to highlight the performance of each according to our criteria list (Carpenter, Hunter & Rheaume, 2008).

1.3.1.2. Conventional Prosthetic legs for Above Knee Amputees

For above knee cases, which are more complex and involve the problem of flexing the knee and coordination between the knee and the foot, we reviewed two low cost options, the Jaipur leg and the (not decided yet) in an attempt to highlight their efficiency compared to their low cost and assess their performance regarding the swing and stance of walking and other essential criteria (Goodier, 2011).

1.3.1.3. Conventional Upper extremity Prosthetics

Our low cost solution line also includes the conventional upper extremity prosthesis which consists essentially of a harness connected to an arm through flexible cables which, with the help of muscles and body movements, serve the function of moving the arm, a ritual that can be viewed as complicated and that requires training, but the only choice available at low cost (Harrison, 2008).

    - Innovative Smart or Intuitively Controlled Prosthetics

Our second solution line is an innovative one and includes major breakthroughs in lower and upper limb prosthetics that are viewed in the West as smart tools that restore normal life.

1.3.2.1. The BiOM Ankle for Below Knee Amputees

For lower limb below knee cases we assessed the BiOM ankle which mimics the actual foot in all details and gives the amputee a natural gait without the additional effort and energy needed to raise a conventional prosthetic foot (U.S. Department of Veteran Affairs, 2012).

1.3.2.2. The BiOM AK for Above Knee Amputees

For above knee prosthetics, we performed a detailed study of the BiOM AK which comprises the BiOM ankle coupled with Otto Bock’s X2 knee, a highly advanced microprocessor knee that thinks for itself. The BiOM AK is a leg that is able to climb stairs with ease, go up and down steep gradients and can do anything a normal leg does without additional thought or energy (U.S. Department of Veteran Affairs, 2012).

1.3.2.3. Myoelectric Arms Coupled with Targeted Muscle Reinnervation

For upper extremity prosthetics, we examined two recent brain controlled prosthetic arms. Our first choice is the myoelectric arm coupled with the Targeted Muscle Reinnervation technique. Myoelectric prosthetics are prosthetics that receive signals from muscles and respond to them by movement achieved by motors in the arm. The Targeted Muscle Reinnervation procedure is a surgical intervention where the nerves from the amputated limb are directed to a nearby muscle and allowed to grow there. These now healthy nerves are able to receive signals from the brain and through electrodes on the skin transmit the signals to the prosthetic arm and order it to move, making the myoelectric arm a major solution for all amputees who have access to it (Layton).

1.3.2.4. The Artificial Muscle Operated Arm

Our second choice of upper limb prosthetics is the Artificial Muscle Operated arm which is a thought operated pneumatic arm that doesn’t require any prior surgery. The device comprises a headpiece that sends signals emanating from the brain to a computer inside the arm which then processes the signals and sends the arm into motion. The arm has no motors and moves when air fills certain artificial muscles while certain valves open or close according to computer commands. The AMO is a very recent design and its light weight and cost vs. efficiency value has attracted many manufacturing companies in the West (Trout, 2011).

  - Purpose and Overview: Study of the Feasibility of the Two Possible Solutions

Given the fact that prosthetics are a medical and human right and that Lebanese landmine and UXO victims deserve having access to the best prosthetic devices to restore their normal lives, and in an attempt to assist these victims out of their low socio-economic status, we tried to  investigate which of our proposed solution lines, the conventional or the advanced, would prove more feasible, taking into consideration the economical situation and lack of funds for these victims and at the same time their urgent need for efficient limbs that can help them up on the economical and social scales. Though at first glance the conventional solution might seem more feasible economically, the highly advanced second solution has proven more efficient in the West and has given amputees a new and an almost normal life. Unfortunately, no similar or clear studies exist concerning Lebanese amputees to determine the efficiency and scope of smart and brain controlled prosthetics in enhancing their quality of life.

In our research we tried to fill in this gap and to assess to what extent innovative prosthetics can change the lives of Lebanese landmine and UXO victims and open new horizons for them, otherwise impossible, along with an attempt to determine the feasibility of such innovative devices in Lebanon.

- Methodology

The method used in this report comprises two parts: The assessment criteria and the procedure.

  - Assessment Criteria

The criteria used to assess the different lines of solutions are divided into two parts: legs and arms.

    - Assessment Criteria for legs

The prosthetic feet and legs are evaluated against the following five criteria:

- Versatility of mobility and associated stability:

This criterion includes the range of activity provided by the prosthesis such as walking, sitting, stair ascent and descent, driving, biking, backward walking, etc, in addition to the different aspects of the foot or leg in terms of articulation and stability.

- Walking speeds and accompanying gait quality:

The Walking speeds and accompanying gait quality criterion determines to what extent the prosthesis accommodates variations in walking speed and how does the increased speed affect gait quality and symmetry.

- Metabolic demands:

The Metabolic demands criterion determines how much muscle work and energy are required on the part of the amputee to move a leg.

- Practicality of usage:

This criterion examines the weight, maintenance needs, training period, durability, water resistance, cosmetic appearance and other relevant features of prosthetic limbs.

- Cost:

The cost criterion highlights the range of cost of the prosthesis and relevant factors.

    - Assessment Criteria for Arms

The prosthetic arms in our report are evaluated against the following five criteria:

- Elbow, wrist and fingers: range of motion, articulation and associated scope of   functionality:

The Elbow, wrist and fingers: range of motion, articulation and associated scope of   functionality criterion mainly deals with the degrees of freedom of the prosthetic arm and the associated degree and efficiency of functionality.

- Control and powering system:

The control and powering criterion deals with the control mechanism of the arm such as mechanical or brain controls and the powering system of the arm which can be body muscle energy, electrical energy and others.

- Weight and comfort factors:

The weight and comfort factors criterion deals with features that are crucial to the usability of a prosthetic arm such as the weight, accompanying restricting or limiting factors and other related features.

- Practicality of usage:

This criterion examines the maintenance needs, training period, durability, water resistance, cosmetic appearance and other relevant features of prosthetic arms.

- Cost:

The cost criterion highlights the range of cost of the prosthesis and relevant factors.

  - Procedure

In order to evaluate the prosthetic limbs mentioned above, we used a two-fold procedure comprising a review of the pertaining literature and a field research.

    - Review of Literature

Our review of the literature includes several steps listed below.

2.2.1.1. Review of Literature on Landmines

We conducted a review of the literature on the landmine situation and problem in Lebanon.

2.2.1.2. Review of Literature on Prosthetics

We conducted a thorough review of the literature on both the conventional and innovative prosthetics under study. Our review included a variety of specialized journal publications in the field of prosthetics, research papers, dissertation papers, an online book, Webinars, on line videos, and on line visits to production sites.

2.2.1.3. Final Review

We conducted a final review of the literature based on the selected assessment criteria.

    - Field Research:

Our field research includes the steps listed below.

2.2.2.1. Analysis of User Comments

We performed an analysis of the user and media comments on the iWalk website, for a user side perspective of the BiOM ankle, in addition to different You Tube videos and online TV interviews with Walter Reed Army Medical Center patients, DEKA test pilots, the designers of the Jaipur knee, the BiOM ankle, the Luke arm and the AMO arm, as well as videos on the Jaipur Foot website.

2.2.2.2. Webinar

We have also analyzed the contents of a webinar on the Otto Bock official site for an analysis of the X2 knee.

2.2.2.3. Interview with Dr. N. Sabah

We conducted an online interview with Dr N. Sabah, an expert in biomedics and a professor at AUB for further clarifications concerning the AMO arm.

- Results

The results of our assessment of the different categories of prosthetics in this report are listed below; however, for a better understanding of our results on the part of the reader, we will start with a general overview of the human gait cycle.

  - An Overview of The Human Gait Cycle

A thorough understanding of the different stages and terminologies involved in the gait cycle is essential before our evaluation of the different proposed lower limb prosthetics.

Lower limbs have three basic functions. They must sustain the weight of the person, enable movement and maintain stability and balance (Prahald, 2003).

The ankle joint which is a hinge type joint, has a high articulation potential, and is of primary importance in walking and in maintaining equilibrium (Prahald, 2003). It provides for the dorsiflexion, plantar flexion, eversion and inversion and supination and pronation of the foot (2003).

The gait cycle comprises two parts: the stance phase and the swing phase, achieved through the alternating movements of the two legs (Prahald, 2003). It begins when the heel of one leg strikes the ground and ends when the same heel strikes the ground again (2003).

The stance period includes the movements of the leg while the foot is still on the ground and the knee is in extension (Prahald, 2003). The swing phase follows with the toe leaving the ground and the accompanying bending of the knee (2003).

The table below provides a summary of the different stages involved in the gait cycle.

Table 1: Stages of the Human Gait Cycle                                                                                                                                                     ©2003, University of Michigan Regents, (Prahald, 2003)

The following  illustration shows the stance and swing phases, along with the different positions of the leg in each phase. Note that the stance phase takes up around 60% of the gait cycle and the swing phase  takes up the remaining 40% (Prahald, 2003). There is a period of overlap in between, also known as the double support, where both legs are on the ground supporting the body weight (Mathur, n.d.). The longer the double support lasts, the slower the walking speed would be (Mathur, n.d.).

Figure 3: The phases of the Human Gait Cycle                                                                                                                                                         ©2003, University of Michigan Regents, (Prahald, 2003)

It is important at this stage, to point to the importance of the different flexions of the foot to achieve successful and normal walking.

Dorsiflexion, as illustrated in the table below, is of primary importance, so that the foot establishes firm contact with the ground heel first, absorbs the ground reaction forces and stores the energy needed for loading in order to allow the foot to clear the ground in swing phase (Prahald, 2003). A normal foot provides 0-20 degrees of dorsiflexion (Lee, & Moroz, 2009).

Plantar flexion, on the other hand and as illustrated in the table below, is also of primary importance, because it is the move that provides the necessary propelling force needed to lift the foot off ground and start the swing phase (Prahald, 2003). The normal foot provides for 0-50 degrees of plantar flexion (Lee, & Moroz, 2009).

Inversion, eversion, supination, pronation, adduction and abduction, as illustrated below, are all due to the hinge type ankle and provide for stability and equilibrium and are of primary importance for walking on uneven ground (Prahald, 2003). Note that the normal foot provides 0-35 degrees of inversion and 0-25 degrees of eversion (Lee, & Moroz, 2009).

The following is a table illustrating the different basic movements of the foot (Prahald, 2003).

Table 2: Basic Foot Movements                                                                                                                                                                       ©2003, University of Michigan Regents, (Prahald, 2003)

A rough summary of the gait cycle beginning with the stance phase and followed by the swing phase is as follows:

Stance: heel strike→ foot flat→ heel off→ knee bend→ toe off→ Swing: Initial swing→ mid-swing → deceleration→ beginning knee extension (Mathur, n.d.).

Finally, it should be noted that although the ankle and the foot are the basic elements in walking, many other body parts and movements are also actively involved to provide movement and stability, such as knee motion, pelvic motion, muscle activity of the leg, hip, abdomen and even the upper body (Mathur, n.d.).

3.2. Below Knee Prosthetics

Below knee prosthetics are prosthetic legs used for below the knee amputees.

3.2.1. General Overview of Prosthetic Feet

Artificial limbs have improved to a great extent from the time of the wooden and clunky peg leg (Gordon, Ardizzone, 1960). Today a wide variety of prosthetic feet exists. These varieties differ in their material, design, flexibility, the amount of shock absorption they provide, the amount of energy storage and restoration they achieve, the type of axis and the associated degree of articulation they possess, and of course the degree of adaptation to individual needs and life style they provide.

Today’s prosthetic feet can be subdivided into conventional or unpowered prosthetics and the recent powered ones (Versluys et al., 2008). The conventional types are also subdivided into low cost, solid ankle and single axis feet, more expensive double and multi- axis energy storing feet and the wide number of the variations of the Flex Foot which is a high cost and highly advanced unpowered foot with substantial energy storing and restoration capabilities (2008). The powered prosthetic feet, on the other hand, are motor-powered to help provide for the lost muscles (2008). Some have sensors and microprocessors, thus they can provide a walk that doesn’t require thinking (2008).

In this report, we will discuss two varieties of low cost prosthetic feet: the SACH foot and the Jaipur foot and one variety of high cost powered feet, the BiOM ankle. Our study entails the analysis of the performance of these prosthetics according to the criteria mentioned above

3.2.2. Conventional Prosthetic Feet

Conventional prosthetic feet are passive or unpowered artificial legs that replace the amputated part of the leg and restore to varying degrees the amputee’s ability to move and walk using the remaining sound muscles of the body.

3.2.2.1. The SACH Foot

Figure 4: © Sheck and Siress, The SACH Foot

SACH is an acronym of solid ankle and cushion heel. The SACH foot was designed and developed in 1955 by Elberhart and Radcliffe in the University of California at Berkeley and it is the most popular and used foot prosthetic to this day (Gailey, 2005). Its basic design is simple and comprises a rigid wooden keel enveloped in rubber and a wedge shaped and rubber compressible cushioned heel (2005). The ankle and the foot are combined into one component, making the ankle solid (2005). Today’s SACH feet come in three varieties of heel stiffness according to need ranging from firm to softer (Gordon, & Ardizzone, 1960). The following are the results obtained from our analysis of the SACH foot according to the criteria mentioned above

Figure 5: The SACH Foot                                                  ©American Academy of Orthotists and Prosthetists

.

3.2.2.1.1. Versatility of mobility and associated Stability:

The compressible heel of the SACH foot acts as a substitute for dorsiflexion and a shock absorbent in the initial phases of the cycle, mainly heel strike and mid-stance (Gordon, & Ardizzone, 1960). The compressible heel also serves in providing for a pseudo-plantar flexion after heel strike (Edelstein, 1988). The rigid keel serves as a stable platform that bears and balances the weight of the amputee, especially in mid-stance (Carpenter, Hunter, Rheaume, 2008).  These characteristics make the SACH foot a preferred choice for low activity and elderly people with little demand for dorsiflexion (2008). Moreover, with its rigidity, the SACH foot provides stability and confidence, especially at low speeds (Marinakis, 2004). The firm cushioned heel variety is especially useful for use as an initial prosthesis for amputees in the first stages of rehabilitation (2004).

The SACH foot with its solid ankle and lack of articulation at the ankle joint combined with the cushioned heel hinders plantar flexion, necessary for loading to start the swing phase (Gailey, 2005). It also provides no transverse articulation for the same reason. Consequently, the SACH foot doesn’t provide for real lateral motion such as eversion and inversion (Gordon, & Ardizzone, 1960). The flexible sole that extends all the way to the front foot, can bend slightly to adjust to uneven ground and absorb ground forces from several angles (1960). This leads us to conclude that the SACH foot is the prosthetic of choice for indoors and level ground walking and its accommodation to uneven ground is only partial and insufficient for long walks on uneven terrain (1960).

SACH feet also pose a difficulty in climbing stairs and depressing the accelerator of cars, making climbing stairs and driving very tiring (Gordon, & Ardizzone, 1960). This is due to the rigidity of the ankle and the lack of dorsiflexion and plantar flexion, leading to additional use of muscles and fatigue (Gailey, 2005).

Therefore, the SACH foot due to its rigidity doesn’t provide much versatility in walking, but its rigidity gives the user stability and confidence.

3.2.2.1.2. Walking Speeds and Associated Gait Quality:

The solid ankle of the SACH foot reduces the range of motion of the ankle and the foot complex leading to lower cadence or walking speeds with an increase in late stance duration (Marinakis, 2004).

The compression of the cushioned heel of the SACH foot, combined with the rigid ankle, tend to increase the time from initial contact with the ground at heel strike to loading response at foot flat to toe off (Gailey, 2005). Rather than promoting plantar flexion, the cushioned heel delays it (2005). The consequence of this delay is an asymmetry between the intact leg and the prosthetic one in gait during walking, in addition to a decrease in self-selected speed, as the sound limb would be in late stance and has to wait for the prosthetic foot to achieve foot flat (2005). This means that the sound foot has to stay on the ground for a longer time than it should normally. This waiting period is twice as much as with biological feet (Edelstein, 1988). This effect further extends to the knee (Burger, n.d.). The solid ankle that hinders plantar flexion leads to a delay between heel strike and the initiation of the bending of the knee at around mid-stance or foot flat. The state of the knee in extension will lead the foot to slap the floor at foot flat (Burger, n.d.). These effects are not highly noticeable at low speeds of walking, but they become more pronounced at higher speeds and lead to fatigue due to the additional efforts of the leg muscles and the hip to compensate for the delays and the gait asymmetry which becomes highly noticeable (Edelstein, 1988). This asymmetry is further enhanced by a noticeable hip displacement on the prosthetic side to compensate for the lack of dorsiflexion (1988).

Therefore, the SACH foot is preferably used for low walking speeds to provide maximum comfort and symmetry in gait and it is not suitable for high speeds.

3.2.2.1.3. Metabolic Demands

If the human body is reduced to its center of gravity which lies in the pelvic region, it is noticed that this center of gravity achieves its highest position at mid-stance and its lowest position in double support (Mathur, n.d.). In walking, the movement of this center is related to energy expenditure that is highest at the highest position and lowest in double support assuming the shape of a sine wave (Mathur, n.d.). To raise this center of gravity, the body uses energy, some of which is stored in the foot during heel strike and loaded in plantar flexion and the other part comes from the muscles involved. (Mathur, n.d.)

In the case of the SACH foot, the energy return efficiency is around 30% and it tends to decrease with increasing walking speeds (Barr et al., 1992). To compensate for the low energy generated during late stance, the hip extensors (muscles) generate greater work from heel strike to mid-stance (Gailey, 2005). Moreover, there is an increased quadriceps and hamstring muscles activity to compensate for the lack of dorsiflexion, along with the displacement of the hip on the prosthetic side, hence the need for more energy (Edelstein, 1988).

3.2.2.1.4. Practicality of Usage:

The SACH foot is a widely used prosthetic foot for several reasons. It is simple and has no movable parts that require regular maintenance, making it a practical foot (Hill, n.d.).  According to the manufacturing site, Otto Bock, it is light in weight ranging from an average of 300g for women to around 530g for large sizes for men, with the exclusion of the adaptors. A variety of sizes for children at different ages also exists. The maximum weight it can (Otto Bock, 2012a) sustain is 125Kg (2012a). It can withstand limited humidity, and contact with water is not very desirable (2012a). It is durable, owing it to the material used, and its average life is around 2.5 years after which the rubber material wears out and may cause loss of balance (Adalarasu, 2011). Although new varieties with natural looking toes have a cosmetic appearance the SACH foot can’t be used barefoot due to fixed heel heights, but it comes with a variety of heel heights especially for women (Ottobock, 2012a).

Finally, the fitting of the SACH foot can take up to 3 months and requires numerous visits to the prosthetist (Jaipur Foot, 2007).

3.2.2.1.5. Cost

The SACH foot is among the lowest in cost in the market today and it is manufactured by several companies around the world including China. The price of the foot depends on the manufacturing site. The ones made in China are the lowest in cost. The European products cost more. The cost, however, does not stop at the foot. The whole prosthetic limb cost depends on the kind of adaptor, the socket and the charges of the prosthetist and can vary from a few hundred dollars to thousands of dollars especially when titanium is used. According to the Jaipur foot site, fitting an amputee with SACH foot can cost up to $8000 (Jaipur foot, 2007). However, the average cost of a SACH foot below knee prosthetic is around $3000 (Blough et al., 2010).

3.2.2.2. The Jaipur Foot

Figure 6: A Sagittal Section of the Jaipur Foot, (Adalarasu, 2011)                                                                                                         ©@ 2011 http://www.ijaest.iserp.org

The Jaipur foot is a handmade lower limb prosthetic (Jaipur Foot, 2007). It was developed in 1968 by Masterji Ram Chandra Sharma, a craftsman, under the guidance of Dr. P. K. Sethi, an Indian orthopedic surgeon, in Jaipur, India (Indian Express, 2008). It is primarily fabricated and fitted by Bhagwen Mahaveer Viklang Shahayata Samti (BMVSS) which is a nongovernmental and nonprofit organization that provides free prosthetics for the poor around the globe (2008). According to the Jaipur website, the BMVSS has already fitted 381 Lebanese landmine and UXO victims in the South of Lebanon (Jaipur Foot, 2007). The Jaipur foot is also used by the International Committee of the Red Cross for similar reasons (Indian express, 2008) and it has been considered the product of choice in developing countries for its low cost and high efficiency (Prahald, 2003).

The Jaipur foot is made of three blocks which makes it very close to the biological structure of the foot (Prahald, 2003). The forefront and heel are made of rubber and the ankle block is made of wood (2003). The whole foot is enclosed in a rubber shell and vulcanized (changed chemically into durable polymer) in a mould to acquire a natural appearance (2003). It has a high degree of articulation and provides for a variety of activities for users (2003).

The following are the results obtained from the analysis of the Jaipur foot according to our criteria.

3.2.2.2.1. Versatility of Mobility and Associated Stability:

Range of motion of the Jaipur foot according to a study performed by Dr. C. K. Prahald (Prahald, 2003):

| Dorsiflexion | 40° |
| --- | --- |
| Plantar Flexion | 0° |
| Inversion | 10° |
| Eversion | 10° |
| Supination | 7° |
| Pronation | 5° |

Table4: Articulations of the Jaipur Foot

The high degrees of dorsiflexion provided by the Jaipur foot lead to a high degree of loading in the stance phase and thus provide for an easy and efficient walk, along with a variety of movements such as driving, climbing stairs and trees, kneeling, crouching, biking, jumping from heights, as well as squatting (Prahald, 2003).

The sufficient inversion and eversion levels provide for stability and enable users to walk with ease on uneven ground and on muddy terrain (Mathur, n. d.)

The transverse rotation of the foot also helps in maintaining equilibrium and gives the users the ability to do all activities done with a normal foot, along with the option of cross legged sitting (Mathur, n.d.).

Figure 7: The Jaipur Leg                                                 ©2007 Bhagwan Mahaveer Viklang Sahayata,

44

It should be noted that the Jaipur foot function is only efficient when used with the Jaipur total contact socket which is made of high density polyethylene, a durable and lightweight material (Jaipur Foot, 2007). The whole prosthetic being made on the spot uses advanced biomechanical alignment techniques for individual fittings, thus making the socket, the axis and the foot in optimum positioning for high efficiency (2007). This allows for an even distribution of weight over the stump and helps give better sensory feedback to the amputee, as well as better proprioception (2007). All these factors, along with the high flexibility and high dorsiflexion degrees contribute to stability in walking and to the comfort of the user, enabling him to adopt an active life (2007).

The Jaipur foot also has a large area which is useful for efficient weight bearing and enhances stability to the degree that almost no initial training is necessary (Kapp, & Cummings, n.d.).

Therefore the Jaipur foot provides versatility of mobility along with a high degree of stability.

3.2.2.2.2. Walking Speeds and Accompanying Gait Quality

Due to the high flexibility in all planes (except for plantar flexion) and the high range of motion of the Jaipur foot, users can vary their walking speeds as desired and can achieve running (Mathur, n.d.). As discussed earlier, this factor is enhanced by the innovative socket structure that distributes the stress rather than localizing it, making long hours of standing and walking comfortable (Mathur, n.d.).

Being a non commercial and nonprofit prosthetic, experiments comparing the gait quality provided by the Jaipur foot to other available ones are rare (Prahald, 2003), but in general all existing literature suggests that the gait quality of amputees using the Jaipur is close to normal especially if compared to the negligible cost of the prosthetic (2003). The Jaipur prosthetic leg accommodates well to all walking speeds and the associated gait can be considered stable and close to normal (Jaipur Foot, 2007). Moreover, the shock absorbing capacity due to the transverse rotation potential of the foot enhances the gait quality by providing comfort at the level of the stump (Sethi, Udawat, & Kasliwal, 1978).

However, after watching several amputees’ gaits at different speeds in various videos, we observed a persistent slight limp and a slight displacement of the hip to the prosthetic side, which is normal in the absence of plantar flexion for proper loading. This observation is confirmed by Dr. Adalarasu in his comparative study of the SACH, the Jaipur and the Madras feet (Adalarasu, 2011).

3.2.2.2.3. Metabolic Demands

The Jaipur foot, being as mentioned a noncommercial product, has not been subjected to extensive testing and comparison experiments in the West (Prahald, 2003). However, in a study conducted by Professor C. K. Prahald, where the Jaipur foot was compared to highly expensive feet like the Variflex which costs $3700, it was concluded that the Jaipur foot, for its low price, matched to a high degree the high cost Variflex in most aspects, including energy storing and release (Prahald, 2003).

The high flexibility of the Jaipur, especially in dorsiflexion, helps the foot absorb the ground forces and store energy for loading (Mathur, n.d.). In addition the resilience of the total contact socket which contributes to an effective distribution of weight and pressure over the stump helps also to distribute the energy efficiently over the whole of the stump (Mathur, n.d.)

Moreover, the transverse rotation of the shank on the foot allows for additional shock absorption at heel strike and enhances the energy storage and redistribution mechanism of the foot (Mathur, n.d). Thus, it reduces additional metabolic demands and in the process reduces fatigue that accompanies or follows additional use of body muscles and energy. It also plays a major role in the redistribution of stress on the stump, making walking more comfortable and saving the energy needed to avoid pain on the stump (Sethi, Udawat, & Kasliwal, 1978, Kapp, & Cummings)

An interesting fact observed by Dr P. K. Sethi about the Jaipur foot is that the sponge rubber used in making the foot acts as a pneumatic joint due to its gradual recovery property after compression   (Sethi, Udawat, & Kasliwal, 1978). This factor enhances the transmission of energy or thrust to the residual leg (1978).

3.2.2.2.4. Practicality of Usage:

The Jaipur foot has a cosmetic appearance and can be used with flat shoes as well as for walking barefoot. It comes in three shades of skin color (Jaipur Foot, 2007).

According to BMVSS, the average weight of a 55Kg person’s lower limb weighs 3.36Kg, while the Jaipur prosthetic weighs only 3.11Kg (Jaipur Foot, 2007). The foot alone is 850g (2007). Recently the BMVSS has signed an agreement with the Michigan Technological University, to improve the structure of the foot and make it lighter and give it an international standard and the new prototype developed weighs only 650g (Sebastian, 2012).

The Jaipur foot is 100% waterproof and can be used in water, mud and all road conditions (Jaipur Foot, 2007).

A factor that adds to its practicality is that it is the only prosthetic that is made on the spot and it only requires around two or three hours of work in total (Jaipur Foot, 2007). However, this fast fitting doesn’t always come without problems, especially given the fact that the prosthetics are made by people who lack proper training, in addition to the hectic pace of work in BMVSS camps (Rosenberg, 2011). Nevertheless, the BMVSS is an organization that constantly works at improving its services and the projects for mass production of the Jaipur foot can reduce this problem.

The Jaipur foot is durable and has a life span of around three years (Adalarasu, 2011).

3.2.2.2.5. Cost

The Jaipur foot is considered by many in the field as the most cost-effective prosthetic foot available. The cost of the foot alone is $5 and the whole prosthetic costs around $30-$35 though it is usually fitted free of charge by the BMVSS (Jaipur Foot, 2007). It’s low cost is due to the fact that it is made of locally available material such as tire rubber and cheap but durable high density polyethylene pipes (Prahald, 2008). According to the BMVSS, future plans for a mass production of the Jaipur foot using polyurethane foam, another cheap, light and durable material are being considered (Jaipur Foot,2007) to reduce the cost by around 40% (Prahald, 2003).

3.2.3. The Power Foot: BiOM Ankle

The Biom is a robotic ankle produced by iWalk (iWalk, 2012). It is one of the most advanced prosthetic feet in the market today. It is the result of the endeavor and persistent work of Dr. Hugh Herr, the founder of iWalk and the director of the Biomechatronics group at MIT Media lab (iWalk, 2012). Being himself a bilateral amputee, Dr. Herr has spared no effort to invent an ankle that mimics and even “emulates” the human biological ankle (Strickland, 2012)).

Figure8: The Power Foot BiOM, (Mancinelli et al., 2011)                                                                                                                                    ©2011 IEEE

The BiOM, as described by Dr. Herr, mimics the biological ankle (iWalk, 2012). A spring substitutes for the lost Achilles tendon and is controlled reflexively by sensors, just like the calf muscles are controlled by the spinal cord, while the lost calf muscle is substituted by a motor (2012).

The BiOM is a sophisticated technological device. It comprises a motor connected to two springs which adjust the stiffness of the foot (Herr, Grabowski, 2011). These are in turn connected to heel and forefront springs which serve to provide elasticity (2011). The motor is a 200W DC one that is powered by a lithium polymer battery in the system and it provides instantaneous power to the foot (2011). The prosthesis also includes sensors that comprise the motor shaft, encoders, three accelerometers and three gyroscopes, the combination of which provides the reflexive response unique to the BiOM (2011).

Figure 9: The structure of BiOM, (Herr, & Grabowski, 2011)                                                                                 ©rspb.royalsocietypublishing.org

The powered neuromuscular model of the BiOM provides the prosthesis with the following (Au et al., 2007):

- A high degree of articulation

- Control of the ankle position in the swing phase

- Control of the stiffness of the foot according to need

- Capability of responding in real time

- Shock tolerance at heel strike to protect the foot from damage

The following are the results obtained from the analysis of the BiOM ankle with respect to our selected criteria.

3.2.3.1. Versatility of Mobility and Associated Stability

The BiOM provides a high degree of controlled articulation in all planes with the addition of powered plantar flexion and thus provides all the necessary factors to achieve a successful and normalized walk (iWalk, 2012).

With its three microprocessors and twelve sensors, it has the ability to adjust the ankle angle, stiffness and damping 500 times/second, giving the user equilibrium, stability and confidence in all conditions of terrains (iWalk, 2012). Moreover, the sensors detect the position and state of the foot at all times and provide it with the ability to react in real time, thus enabling it to adapt to all types of terrains (Prokaza, 2012).

The motor provides the ankle with power in each step, and peak power at late stance or plantar flexion and thus propels the user forward, giving him an effortless and normalized mobility (iWalk, 2012).

According to our analysis of the BiOM user’s comments on the iWalk site, users are able to walk on rough or uneven terrain, climb and descend stairs, climb and descend a ramp and achieve the same range of activities as with a biological foot without even having to look down with each step (iWalk, 2012). As mentioned, the foot propels them forward and the sensors measure the position of the foot in space and determine the degree of flexion needed (Strickland, 2012).

The high degree of mobility provided by the BiOM powered foot is also associated with stability. The stability is particularly achieved by the fact that the foot behaves like a spring with variable stiffness beginning with heel strike until foot flat or mid-stance, after which it provides additional power at plantar flexion until toe off or the beginning of the swing phase (Au et al., 2007). During the swing phase, the system orients and controls the foot according to need (2007). In addition, the capability of the foot to react in real time improves and even normalizes symmetry (Prokaza, 2012). This means that the user doesn’t have to wait for the foot to react and the reaction occurs naturally without effort or thinking (2012). The combination of these behaviors provide the amputee with stability and most importantly, confidence, as most users have reported forgetting their disability or even the prosthesis being there (iWalk, 2012).

Moreover, the normalized range of activity provided by the BiOM ankle is accompanied by health benefits, since increased activity reduces overweight problems and related ailments such as cardiovascular problems (Knapp, 2012).

In short, the BiOM ankle’s smart technology provides users with the range of activities of a normal foot and gives them a normalized degree of stability in walking.

3.2.3.2Walking Speeds and Accompanying Gait Quality

Graph 1:Preferred Velocity         ©rspb.royalsocietypublishing.org, (Herr, Grabowski, 2011)

According to Dr. Herr, the BiOM responds reflexively to all walking speeds and helps push the foot off the ground, just like biological calf muscles do (Eilenberg, Geyer, & Herr, 2010). He adds that the more the user pushes on the foot, the more the foot will push back (Strickland, 2012), and that the power provided by the ankle increases with increasing speeds and vice versa (Eilenberg, Geyer, Herr, & 2010). Thus, the BiOM adapts to all walking speeds, leading to higher self selected speeds when compared to other passive prosthetic feet (Au et al., 2007).

Gait quality is directly related to the net work produced by the foot (Mancinelli, 2011).The energy provided by the BiOM at toe off serves the function of improving user gait quality. Moreover, the plantar propelling force generated by the motor decreases the need for additional use of leg and hip muscles, as is the case with unpowered prosthetics where an obvious hip displacement is observed as a result of the extension of the hip muscles to compensate for the low energy generated at late stance. The consequence of this extra power is more symmetry in gait and even a normalized gait (Au, & Herr, 2008).

The symmetry in gait is also enhanced by the personal tuning option of the BiOM, to adjust it to individual gait cycle needs. It can be tuned over Bluetooth using an android application (Prokaza, 2012). In this way, the various parameters of the ankle are adapted according to user preference, such as stiffness which is responsible for determining the time

Figure 10: The Android Application for Tuning the BiOM, adapted from http://www.knoxnews.com/news/2012/aug/23/leaps-and-bounds-marine-says-new-bionic-foot-get/

required by the foot to descend and touch the ground (Strickland, 2012). The individual tuning feature is particularly essential for achieving gait symmetry for unilateral amputees in order to match the prosthetic leg which is superior in action, to the biological one (Prokaza, 2012).

From the preceding information, we can conclude that the BiOM ankle provides a consistent, natural and symmetric gait at variable speeds.

3.2.3.3. Metabolic Demands

Depending on the level and cause of amputation, as well as the selected walking velocities, amputees, in general, spend around 10%-60% more energy in walking than non-amputees, due to their need to use other body muscles to compensate for the lost ones in the leg (Au, & Herr, 2008). The BiOM, being a power supplying ankle, provides the amputee with a propelling force with every plantar flexion, thus decreasing the metabolic demands of walking.

Graph 2:Metabolic Cost of Transport ©rspb.royalsocietypublishing.org

Passive elastic foot  ̶

BiOM ̶

Biological foot  ̶

The efficiency of the BiOM being around 67%, it provides 30J of energy or positive work with every step, 20J being produced in the stance phase. This is the same amount of energy required for a person with sound legs weighing 80Kg and walking at a speed of 1.75m/s (Herr, & Grabowski, 2011). This energy boosts every step and pushes the user forward. As mentioned earlier, the power produced increases with increasing speeds and decreases with decreasing speeds, sparing the user additional efforts and energy. Moreover, this energy is produced by the motors and not by body muscles. The result is a decrease in the metabolic cost of ambulation and hence a decrease in fatigue associated to walking with passive prosthetics, in addition to the preservation of the sound parts of the body (2011).

The results of a research done by Dr. Hugh Herr and Dr. Alena Grabowski, concerning metabolic demands associated with the use of the BiOM, a passive energy storing foot, and a sound foot, show a decrease in the metabolic demands associated with the BiOM as shown in the adjacent graph. The curves show that at speeds less or equal than 1m/s, the metabolic demands of the BiOM users are equivalent to the demands of non-amputees which is a feature unique to the BiOM (Herr, & Grabowski, 2011).

In another study, where the BiOM was compared to the Ceterus which is a rather advanced and unpowered energy storing and releasing foot, the results showed a 54% increase in the power generated by the BiOM in late stance, along with a decrease of 8.4% in the oxygen consumption of users (Mancinelli et al., 2011). The same study found that the overall power provided by the BiOM was twice that provided by the passive energy storing foot, the Ceterus (2011).

All these results lead us to conclude that the BiOM, through its capacity of supplying a high amount of power, decreases and even normalizes the metabolic demands of walking for amputees.

3.2.3.4. Practicality of Usage

The BiOM weighs 2Kg, which is the weight of the foot of an 80Kg person (Herr, & Grabowski, 2011). This close to normal weight makes it a comfortable device to wear. Moreover, this weight decreases during walking due to the propelling upward force provided by the motor, which even cancels the weight of the prosthetic (Herr, & Grabowski, 2011).

The battery that powers the foot weighs 0.22Kg (Herr, & Grabowski, 2011). It is a rechargeable lithium polymer battery that lasts for about one day of usage (2011). It can provide for 4000-5000 steps, equivalent to 4Km-5Km at a velocity of 1.7m/s (2011). The normal average is 3060 ± 1890 steps per day for active people (2011).

The BiOM comes with a package of three batteries and a dual battery charger (Fierce Medical Devices, 2012). One battery is used for walking, the other for back-up and the last for recharge (2012). The charging of the batteries requires around an hour (iWalk, 2012). The new batteries have a 20% increased battery life (Fierce Medical Devices, 2012).

It also provides an intuitive user interface which includes LED lights and a low battery buzz to inform the user of the charge status of the battery (Fierce Medical Devices, 2012).

The BiOM doesn’t require long training periods for adaptation. It takes a healthy amputee less than 20 minutes to adapt to the foot (Au et al., 2007).

The BiOM is an electronic device and can’t withstand moisture, dust and dirt, as these can damage the system (iWalk, 2012).

It is suitable for amputee weights of 170-250lbs (iWalk, 2012)

The functioning BiOM ankle is associated with a bit of noise (Strickland, 2012).

3.2.3.5. Cost

The cost of the BiOM ankle is within the range of $40000-$60000, the average cost being $50000 (Wynn, 2012).

3.3. Above Knee Prosthetic Limbs

An above knee prosthesis is a prosthetic limb that replaces a lost leg for transfemoral amputees, that is for amputees who have lost their knee and parts of their thigh. Above knee prosthetics are more complex than below knee foot ankle and shank units and amputees face major difficulties in regaining mobility (Mc Gimpsey, & Bradford, n.d.). These difficulties are due to the loss of the support provided by the long bones of the leg and the lost knee and ankle joints and they vary according to the length of the residual limb, the result being an inability to stand or walk freely. (Mc Gimpsey, & Bradford, n.d.) In clearer terms, the main complexity lies in the movement of the knee and the associated movement of the foot complex (Wilson, 1968), in addition to the fact that the stump needs to accommodate the body weight with every step taken (Mc Gimpsey, & Bradford, n.d.). Moreover, the whole motion of the leg is controlled by the amputee, particularly by the residual thigh muscles. Consequently, the amputee can control only the knee part of the prosthetic and he is unable to control the shank, ankle and foot (Wallach, & Sabiel, 1970). In today’s efficient AK prosthetics, the role of the knee extends to that of control of the motion of the below knee parts.

According to Dr. Wallach, an efficient AK prosthesis must have a normal appearance, permit safe and comfortable mobility at normal rates without physical and mental effort and has to provide a normal gait (Wallach, & Sabiel, 1970). These features have been a challenge for prosthetists and workers in the field and much research has been dedicated to improve AK prosthetics and especially prosthetic knees, leading to more than 100 knee types of knees (Dupes, 2008). Before specifying the types of knees and their evaluation according to our criteria, a brief explanation and overview of the components of AK prosthetics, their types, and their way of functioning is necessary for a better understanding of our assessments.

3.3.1. AK Prosthetics: Components, Parts and Functioning modes

AK prosthetics are made of 4 essential components: The socket, the knee, the shin and the ankle-foot unit (Wilson, 1968).

Figure 11: The Above knee Prosthesis                                                                                                                        ©http://old.limblossinformationcentre.com

a. The Socket:

The Socket is of primary importance in AK amputees, because it bears the whole body weight at a certain stage in walking. Sockets have improved greatly and today’s total contact sockets have proven highly efficient in preventing oedema of the stump and can distribute the body weight evenly over the whole of the stump, for better comfort and less pain (Mc Gimpsey, & Bradford, n.d)  . In some cases a pelvic belt or shoulder straps are added, especially for short stumps or weak patients (Wilson, 1968).

b. The Knee Unit:

The knee is the main differentiating factor in AK prosthetics. It provides stability, safety against buckling, and control of the lower part of the prosthetic (Dupes, 2008). Prosthetic knees can be divided into two types: Mechanical and computerized (2008).

i. Mechanical knees are further subdivided into single axis knees and polycentric knees (Dupes, 2008).

- Single axis knee: It is a simple hinge like joint that can only flex and extend in one plane (2008).

- Polycentric knee: It has multiple axis of rotation and is closer to the biological knee in action (2008).

ii.Microprocessor Knees: These are recent technological devices that have sensors that detect movement and microprocessors that provide for adjusting fluid control in the knee, and can thus provide a natural and safe gait for amputees (Dupes, 2008).

c. Control options for Mechanical knee units:

All prosthetic knees need motion control systems, especially in the swing phase. The reason for this is that in late stance, the thigh is forced forward by the thigh muscles of the amputee and the knee starts flexing (Mc Gimpsey, & Bradford, n.d). As a result, the foot leaves the ground and the shin, together with the foot are forced into a backward rotation as the heel rises above the ground (Mc Gimpsey, & Bradford, n.d). If this rotation is allowed to take place without control or resistance, the heel might rise too high and lead to gait complications (Mc Gimpsey, & Bradford, n.d).

©http://www.oandplibrary.org/al/1968_02_001.asp                                                                                                                                   Figure 12: The motion of the lower leg in swing phase

Several control mechanisms exist. The constant friction knee, for example, has a fixed and adjustable resistance mechanism that prevents buckling, but provides a single speed of walking (Dupes, 2008). The variable friction knee adjusts the resistance to knee angles and walking speeds and provides for a variety of speeds (Wilson, 1968). Higher cost, fluid controlled knees provide more efficient control of the swing phase and provide even a smoother and safe walk (Wilson, 1968).

Some knees also provide for stance control which provides for more stability in standing. A low cost variation is the manual lock which enables the amputee to lock the knee whenever desired, such as when standing or walking on rough terrain (Dupes, 2008). More advanced knees provide for weight controlled stance control, where the increase in weight on the shank initiates a braking system to prevent buckling (Wilson, 1968).

d. Foot and Ankle Unit:

Most prosthetic systems can accommodate any type of foot-ankle units which can be chosen according to age, need, affordability and other factors (Wilson, 1968).

3.3.2. Low Cost Conventional or Mechanical AK Prosthetics

From the previous discussion, we can conclude that the differentiating component of AK prosthetics is the knee unit and consequently our analysis of AK prosthetics will be focused mainly on the knee. The low conventional prosthetics examined in this report are the single axis, constant friction mechanical knee and the polycentric Jaipur knee.

3.3.2.1. Single Axis Constant Friction Knee

© http://www.amputee-coalition.org/military-instep/knees                                                                                                                                Figure 13: A Single Axis Knee Unit

The single axis knee is the most basic and simplest type of prosthetic knees and it is the most commonly used prosthetic knee in developing countries where because of poverty, rigid knee prosthetics are still in use. It is also one of the least expensive, lightest in weight and one of the most durable flexible knees (Dupes, 2008).

The single axis knee resembles a door hinge in its mechanism and it flexes freely at the level of the knee (Michael, n.d.). It is usually associated with a constant friction providing unit for some degree of control of the shank during swing (Dupes, 2008). It provides no stance control and it mainly depends on the amputee muscle powers for stability (Michael, n.d.). A manual lock option can be added for additional safety (Dupes, 2008).

3.3.2.1.1. Versatility of Mobility and Asso

Figure 14: The Single Axis Knee Prosthetic Leg ©http://img.medscape.com

ciated Stability:

The single axis knee-shin complex operates like a pendulum (Wilson, 1968). It provides for simple forward walking with the swing motion of the leg controlled by the constant friction in the knee. It is actuated by the power supplied by the residual limb and the hip (1968). Therefore, it needs strong amputee muscles. Moreover, it lacks stance control, making buckling of the knee a serious problem in the stance phase and in standing (Michael, n.d.). The control and stability depend on the amputee muscles (Michael, n.d.). It can be used to walk on even surfaces like indoors wooden floors, but walking on a carpet or on grassy land, or carrying a rather heavy weight such as a baby, is a challenge for the weak amputee (Michael, n.d.). That is why, an additional lock mechanism is added for some amputees that locks the knee in extension for standing and challenging surfaces in order to avoid voluntary muscle control and fatigue (Dupes, 2008). However, walking with an inflexible knee is also difficult and awkward, but can provide more stability, especially for weak amputees (Dupes, 2008).

Sitting with a flexed leg is possible due to the flexibility of the knee (Michael, n.d.). Manipulating stairs is also possible for strong amputees (Michael, n.d.).

Therefore, the single axis knee provides for a limited range of mobility and the associated stability depends on the amputee’s ability to control the knee, unless it is manually locked.

3.3.2.1.2. Walking Speeds and Accompanying Gait Quality

The constant friction in the single axis knee imposes a constant speed of walking (Michael, n.d.). However, this friction is adjustable to accommodate individual walking speed preferences (Michael, n.d.). The response of the knee to variable speeds of walking is linear, meaning that it moves at a constant rate (Greene, 1983).

The gait pattern accompanying the single axis constant friction knee has been extensively studied. According to Dr. B. Farber and Dr. J. Jacobson, it is a rigid gait with an accompanying impact on the pelvis, “abnormal kinematics” and a difficulty in increasing walking speed (Farber, & Jacobson, 1995). Moreover, the prosthetic leg needs to be shorter than the normal leg, since it extends during walking and acquires a length exceeding that of the sound limb (1995).

In another detailed study concerning the gait quality of the single axis knee by Murray et al. the results reveal a highly asymmetric gait, as shown in the table below (Murray et al., 1980).

©http://www.rehab.research.va.gov/jour/80/ 17/2/murray.pdf                                                                                                Table 4: Comparison of prosthetic and normal leg gait features (Murray et al., 1980)

The following is a summary of the gait characteristics and the impact of the single axis knee on overall symmetry during ambulation as listed by Murray et al.:

1-Walking speeds deviate from normal because of longer than normal gait cycles and abnormally wide uneven stride lengths (Murray et al., 1980). Amputee walking speeds were found to be 2/3 the speeds of normal people (1980). The reason is that the stance phase on the prosthetic side was longer than on the normal side while the swing phase on the sound size was shorter (1980). Moreover, with biological legs, strides were approximately equal in length, while the abnormally wide strides on the prosthetic side were accompanied by 8degrees of out-toeing on the sound leg side (1980).

2-There was a pronounced lateral movement of the body to compensate for stability during prosthetic limb single limb support, in addition to forward and vertical movements that added to the asymmetry of the gait (Murray et al., 1980). The head was noticed to deviate more to the prosthetic side in stance and the neck and head underwent a vertical forward motion with every step forward (1980). The shoulders showed excessive flexion on the sound side whereas the arms exhibited excessive extension at the elbow (1980).

3-The hip on the prosthetic side exhibited an abrupt and far from smooth transition between flexion and extension, accompanied by hip- hiking, characteristic of most conventional prosthetics (Murray et al., 1980).

4- Other particularities observed were an excessive rise of the heel with an excessive toe clearance on the prosthetic side showing a low proprioception and a premature plantar flexion on the sound limb side, possibly due to the inability of the prosthetic knee to flex at the right time in early stance (Murray et al., 1980).

As Murray and al. explain, these gait abnormalities are the only way to ensure safe walking. They serve to compensate for the deficiencies of the prosthetic leg using the single axis constant friction knee (Murray et al., 1980).

In short, the single axis constant friction knee can provide a constant walking speed with a pronounced asymmetry in gait quality.

3.3.2.1.3. Metabolic Demands:

As discussed above, walking with a single axis knee leads to an asymmetric gait associated with multiple movements of different body parts and in multiple directions to ensure safety ((Murray et al., 1980). The result is a high metabolic cost, as these muscles spend energy to operate. Even standing puts a higher than normal metabolic demand on the amputee, unless a manual lock is available. The energy cost of the single axis knee can go up to 80% more than sound knees of non-amputees (Mc Gimpsey, & Bradford, n.d.).

3.3.2.1.4. Practicality of Usage:

The single axis constant friction knee is a simple device; hence it doesn’t require frequent visits to the prosthetist for maintenance of parts (Michael, n.d.). It is ideal for children who outgrow their prosthesis and need to change it frequently. It is also suitable for people living in rural areas with little facilities to visit their prosthetist frequently (Michael, n.d.).

The optional manual lock can prove also practical for users.

It is only suitable for active and healthy amputee and can’t be used by the elderly or weak.

3.3.2.1.5. Cost:

The single axis knee is one of the least expensive types of prosthetic knee. It costs around $100-$500 depending on the production site and the existence or lack of a lock mechanism. However the whole prosthetic including the prosthetist cost, the socket (endoskeletal), shank, ankle and foot (SACH foot) cost around $2000- $3000 (Orthotic and Prosthetic Appliances: Billing Codes and Reimbursement Rates-Prosthetics, 2012).

3.3.2.2. The Jaipur Knee

Figure15: The Jaipur Knee ©http://transition.usaid.gov/scitech/jk.html

The Jaipur Knee was designed in 2009 at Stanford University, by Joel Sander and three fellow students, in collaboration with the Jaipur Foot organization, BMVSS, in answer to the need for a low cost AK prosthetic for developing countries (D-Rev Re:Motion Designs), since 80% of the world’s amputees live in the developing world (USAID Science and Technology, 2010). It is a polycentric knee unit (2010), made of five pieces of plastic, four nuts and a bolt (What Gives 365, 2010). It has a simple design which allows for easy mass production (USAID Science and Technology, 2010). It is made of a low-cost, self lubricating and durable polymer (2010). The Jaipur knee has been selected by Time magazine to be one of the 50 most important inventions in 2009 (What Gives 365, 2010).

3.3.2.2.1. Versatility of Mobility and Associated Stability:

The Jaipur knee is a polycentric knee (USAID Science and Technology, 2010), meaning it has multiple axis of rotation, resembling the biological knee in articulation (Dupes 2008) .

A polycentric knee, in general, is a four bar linkage knee that provides for a non linear or more exactly, a rotational movement of the center of the knee (Greene, 1983). This center changes location according to the angle between the axis of the knee and the shank and the acceleration and deceleration of the shank depend on the position of this center (1983). In simple terms, just like a clock pendulum, when the applied weight decreases, the pendulum slows down and when it is increased, the pendulum swings faster (1983). When the knee flexes and leaves the ground the weight applied by the body decreases gradually, so the backward swing of the shank decelerates (1983). The instantaneous change in the position of the center with respect to the shank controls the acceleration and deceleration of the shank (1983). These actions are timed to occur at certain positions and result in an efficient swing and eliminate the need for mechanical friction used in other types of prosthetic knees (1983). The polycentric knee can thus provide for an efficient swing control (1983).

©http://www.oandplibrary.org/poi/1977_03_ 146.asp                                                                                                                                 Figure 16: The rotational trajectory of the center of the polycentric knee (Radcliffe, 1977)

Hence, the Jaipur knee being a polycentric knee provides for swing control and a successful forward motion of the user. Through the elevated center of rotation of the knee in full extension, it also helps the amputee in exercising control on the knee using his own muscles and achieving successful foot clearance (Radcliffe, 1977).

Another feature of the Jaipur knee as illustrated by one of its designers, Joel Sander, is the ability to provide stance control by gently locking the knee when the leg is in extension (The Jaipur Knee Design Story, Vimeo). This feature provides anti-buckling safety and stability in standing and walking. The stability in late stance is followed by a smooth and easy flexion of the knee as the leg starts swing phase (Michael, n.d.).

All these features make the polycentric Jaipur knee safe and suitable for meeting the needs of moderately active and independent amputees, as well as bilateral amputees (Michael, n.d.). It enables amputees to regain their ability to walk and even use bicycles, an activity barely possible with advanced $10000 prosthetic knees (The Jaipur Knee Design Story, Vimeo). The Jaipur knee also provides stability on uneven terrain (Davies, & Starr, 2009).

Another feature characteristic of polycentric knees is the shortening of the shin length in sitting position, just like the biological knee and leg (Greene, 1983). This feature leads to a more normal appearance in sitting due to the resulting equal lengths of the sound and prosthetic sides (1983). Note that ordinary prosthetic knees don’t provide for this and the prosthetic leg has to be made shorter to achieve this equal appearance.

The Jaipur knee has a high range of motion that can reach 165° (D-Rev Re: Motion Designs). This high flexibility allows for kneeling and squatting (D-Rev Re: Motion Designs).

Before reaching our conclusion, it should be noted that a proper alignment and an efficient and comfortable socket, preferably an ischial containment total contact socket, are crucial for efficient results (Wilson, 1968).

Therefore, the Jaipur knee is a safe prosthetic knee capable of restoring amputee mobility to a reasonable degree. This mobility is associated with sufficient stability as confirmed by 95% of users who have reported the no failure or no buckling feature of the Jaipur knee (The Jaipur Knee Project II-Scaling up the Business, 2012).

3.3.2.2.2. Walking Speeds and Accompanying Gait Quality:

Due to the lack of literature concerning walking speeds provided by the Jaipur knee, our analysis depends on the general characteristics of polycentric knees in general.

According to Dr. M. Greene, the action of polycentric knees is only rotational with respect to the angle between the knee axis and the shin (Greene, 1983). The response of the polycentric knee to walking speed is linear. In simple terms, it can’t accommodate changes in walking speeds in the gait cycle (1983). Therefore, the polycentric knee can only provide for a rather constant walking speed. This fact applies to the Jaipur knee and it is evident in the short films featuring an amputee walking with the Jaipur leg (The Jaipur Knee Design Story, Vimeo).

The gait quality provided by the Jaipur knee is more than acceptable, given its low cost. As mentioned before the rotation of the center of the knee aids the amputee in controlling the leg (Radcliffe, 1977). The elevated center of the knee in extension also relieves the amputee from additional effort and energy expenditure (Greene, 1983). Moreover, the stable stance phase and the smooth transition from stance phase to foot swing, in addition to the controlled swing contribute to a close to normal gait (1983). All these factors combined help the amputee to walk with an acceptable gait with decreased hip-hiking and pelvic forward thrust (1983).

http://whatgives365.wordpress.com                                                                                                                                                                       Figure 17:The Jaipur Knee Gait

However, as with all mechanical prosthetics, there is a short delay between the sound leg and the prosthetic leg in the gait cycle. This is due to the longer time spent by the prosthetic leg in stance phase (Farber, & Jacobson, 1995). Another factor to be considered is that all conventional prosthetics depend on the amputee’s muscle powers and gait quality might vary according to amputee strength of muscles (Michael, n.d.).

To conclude, the Jaipur knee provides for a constant speed ambulation accompanied with an improved gait quality when compared to other conventional prosthetic knees.

3.3.2.2.3. Metabolic Demand:

According to Dr. Murray, only a symmetrical gait can decrease metabolic demands for amputees, because it leads to a low amplitude sinusoidal motion of the center of gravity of the body (Murray et al., 1980). The lower the amplitude of this wave, the lower the energy the amputee needs to spend in walking (1980). This means that the acceleration and deceleration forces need to be reduced or more precisely, controlled by the prosthetic device (1980).

In the case of polycentric knees, as explained by Dr Greene, the energy required during extension in early stance is less than in some other prosthetics like the single axis knee (Greene, 1983). Moreover, and always according to Dr Greene, as the knee shortens as it goes from extension to flexion, the energy losses due to gait defects such as those associated with abnormal hip or pelvic movements are reduced (1983). The shortening of the knee also raises the center of gravity of the body automatically and without effort, hence another contribution to reduced metabolic demand (1983). Moreover, acceleration and deceleration of the leg being controlled by the instantaneous positions of the knee center and in exact timing, leads to precision in the gait and results in efficiency, thus reducing metabolic demands (1983).

The Jaipur knee, being a polycentric four bar knee, follows the patterns described above. We can thus conclude that it has a reduced metabolic demand, just like all polycentric four bar knees, with the addition of the light weight of the Jaipur knee which further reduces energy demands.

3.3.2.2.4. Practicality of Usage:

Being made of a self lubricating and high density polymer, the Jaipur knee, unlike all other polycentric knees made of different metals, requires no maintenance (USAID Science and Technology, 2010). It lubricates itself with each step (2010). It can also be used in rough conditions (2010). The polymer is highly durable and gives the Jaipur knee a long life. The Jaipur knee has been benchmarked to have 3-5 years of life (D-Rev Re:Motion Designs).

The whole prosthetic leg, as with all Jaipur prosthetics fitted by the BMVSS, can be produced on the spot and requires only a few hours to assemble (Jaipur Foot, 2007), owing it to the simplicity of the prosthetic parts and to the fact that the Knee can be fitted on a variety of prosthetic systems (D-Rev Re:Motion Designs) . An amputee might be fitted with an AK leg, given training and sent home in one day, which is impossible with other prosthetics (The Jaipur Knee Project I-Getting the Need Right, 2012) .

3.3.2.2.5. Cost:

The Jaipur knee is the only polycentric design that comes at a low manufacturing cost of around $20, which is much less than the cost of other polycentric knee joints with similar characteristics and which can cost $10000 (The Jaipur Knee Project I-Getting the Need Right, 2012) . The factor behind the low cost is the used polymer that replaces high cost metals like titanium and the simplicity of design that makes it easy to manufacture in mass quantities (USAID Science and Technology, 2010). It is an ideal choice for non-profit organizations like the BMVSS to fit amputees of the developing world free of charge. However, it should be noted that the $20 is not the cost of the whole prosthetic leg, and it doesn’t include the cost of the socket, the pylon and the foot- ankle complex.

3.3.3. The BiOM AK: The X2 Bionic Knee

The BiOM AK is a recent technological advance in above knee prosthetics and it combines the BiOM power foot with a very sophisticated microprocessor controlled knee, the X2(U.S. Department of Veteran Affairs, 2012). It is the outcome of extensive research funded by the Department of Veterans Affairs to enable war veterans restore normal ambulation and even go back to service (2012). The X2 is especially designed for military use, but a scaled down version, the Genium, also exists for civilian use (Martin, 2011).

©http://www.va.gov/health/images/20120719a1.jpg                                                                                                                                       Figure 18: The BiOM AK

The X2 and the Genium comprise seven sensors, a gyroscope, an accelerometer, microprocessors (Edwards, 2011) and a hydraulic knee system (Tan, 2012). They are characterized by automatic and intelligent swing and stance controls, high stability controls, and the ability to react instantaneously to multiple and complex environmental inputs to produce natural motion (Edwards, 2011).

Since the BiOM ankle was discussed extensively in the below knee prosthetics section, the following sections will focus mainly on the X2 or the Genium which has the same basic features as the X2, and assess the performance of such advanced knees according to our criteria list.

3.3.3.1. Versatility of Mobility and Associated Stability:

The X2, as the Genium, comprises a knee angle sensor, a knee moment sensor, an ankle moment sensor, a shank inclination sensor, an axial load sensor, a ground force sensor, an accelerometer that detects inputs in two directions and a gyroscope that detects the position of the leg in space (Edwards, 2011). It has a sampling rate of 100Hz, meaning that it detects changes in the environment and samples the input data at a rate of 100 times in one second (2011). The input components work together and provide instantaneous input to the microprocessor system of the knee which in turn controls the hydraulic actuator to produce output and result in motion by adjusting stiffness and position (Martin, Pollock & Hettinger, 2010). The sensor input system resembles the nerve impulse received in a biological limb, the microprocessor plays the role of the brain and the knee output is analogous to the motor movement of muscles (2010). As in the physical body, this cycle is intuitive, meaning it occurs without mental effort and the knee being an intelligent system, does all the thinking (Orange Coast Prosthetics Inc., 2010-2012). The amputee’s participation is confined to moving his leg by using his residual thigh muscles and/or hip extensors (Edwards, 2011).

From the above specifications of the X2, it follows that the X2 provides for a wide range of activities. The intelligent system is the first in its kind to allow step- over- step stair ascent and descent (Edwards, 2011). It also allows for successful obstacle crossing (2011). This is due to the unique feature of the X2, mainly its capacity to lock the knee in bent position (2011). When an amputee hits an obstacle, the knee blocks within five seconds and instead of stumbling and falling or walking around the obstacle, the amputee can raise his locked and flexed knee above the obstacle and cross it like a non-amputee (2011). For example, if the toe is caught on a carpet, a mechanical knee will keep

Figure 19: Walter Reed X2 Knee Test ©http://www.armytimes.com/news/2012/03/army-amputations-reach-war-time-high-031212w/

swinging and result in a fall, whereas the X2 will block the knee even if flexed, and prevent falling (Martin, Pollock & Hettinger, 2010) .

The X2 has an efficient stance control that allows the amputee to tolerate long standing durations without putting strain on the sound leg (Edwards, 2011). This is possible because flexed knee locking is achieved, as mentioned within five seconds of stationary position of the leg and the amputee can stand on the prosthetic leg as well as the sound leg (2011). The intelligent stance control also provides for smooth and safe walking on uneven and rough terrain, without the risk of buckling (Progressive Orthotics and Prosthetics, 2012).

Moreover, the X2 has a sitting recognition function that is initiated within five seconds of horizontal thigh position (Edwards, 2011). The knee unlocks to allow free movement of the leg, particularly useful in moving in and out of a car or to move the legs out of the way in a movie theater (2011).

The X2 is also unique in that it provides for backward walking (Edwards, 2011). The reason is that it has a swift and controlled swing phase initiation that is independent of the changes of body weight at the knee level (2011). The swing control and the direction of swing depend on the ground reaction forces detected by the sensors and the gyroscope which send signals to the microprocessor and allow reaction in real time (2011). Therefore, while backward walking, the knee doesn’t lock as with other prosthetic knees. On the contrary it adjusts resistance and stiffness like in forward motion and thus allows for a successful backward motion (2011). Moreover, the whole swing process, whether forward or backward, requires no mental effort on the part of the amputee. The knee adjusts automatically (Progressive Orthotics and Prosthetics, 2012).

The versatility of mobility of the X2 is further enhanced by its five modes of function (Edwards, 2011). While one mode can be adjusted and tuned to everyday life activities, the other modes can be adjusted to meet other needs such as biking (2011). The modes are programmed using special software via Bluetooth and activated either by remote control or the tapping of the foot on the ground three times (2011).

The X2 is a passive device and hence, it doesn’t provide for energy needed to move, but it provides the same range of activities and mobility as the biological knee (Edwards, 2011). However, when combined to the BiOM power-foot which provides a power boost with every step (iWalk, 2012), the whole system produces an effortless motion (OPC1media, 2012) and enables the amputee to achieve all desired activities and without additional fatigue or discomfort at the level of the stump, hip, back, or sound leg (Ostrovsky, 2009).

To conclude, the BiOM AK, a combination of the BiOM ankle and the X2 knee, can be considered to be similar to the biological leg in the range of activities, and the stability provided by the system in all stages of the gait cycle can be considered highly efficient.

3.3.3.2. Walking Speeds and Associated Gait Quality

The auto-adaptive stance and swing controls associated with the hydraulic system of the X2 knee provide for the variation of speed in walking (Orange Coast Prosthetics Inc., 2010-2012). An amputee can change from slow walking to running, without having to shift into another mode (Martin, 2011). The whole process is provided by the intelligent system that reacts to forces and adjusts the swing or the bending of the knee in real time (2011).

The associated gait quality is very close to normal, especially for the combination of the X2 and the BiOM ankle (OPC1media, 2012). The BiOM AK provides a high degree of symmetry in the gait of unilateral amputees and the X2 provides a reasonably normal gait for bilateral ones (Eyewitness News 5, 2012). This symmetry accounts for comfort and reduced pain at the level of the sound knee, the back and the hip, as reported by the Walter Reed Patients Army Medical Center (Ostrovsky, 2009). The smart swing and stance controls and the ability to react in real time are behind the symmetrical gait of this highly advanced leg.

The symmetry in gait is further enhanced by the personal tuning options of the knee and the ankle (Edwards, 2011). Moreover, the knee has an integrated alignment tool that provides a virtual display of ground reaction forces and gives recommendations for efficient alignment, thus enhancing output and symmetry in gait (2011).

In short, the X2 can adapt to variations in cadence and when combined with the BiOM ankle can provide an efficient and symmetric gait.

3.3.3.3. Metabolic Demands

The X2 bionic knee is a passive device that doesn’t provide energy, but it enhances even force distribution over the whole leg because of its hydraulic system, thus enhancing energy efficiency (Progressive Orthotics and Prosthetics, 2012). It also absorbs ground forces due to its highly developed stance control and flexion control (2012). In addition, its ability to adjust the knee resistance reduces the amputee’s efforts and need to control the knee with excessive muscle force, as in conventional prosthetic knees (2012). Thus, it reduces energy demands and fatigue (U.S. Department of Veteran Affairs, 2012). This energy efficiency is accompanied by a reduced psychological strain, since the amputee doesn’t need to think about every step he takes and how to control it (Hafner, 2007). This leads to an increased confidence and enables the amputee to achieve multi-tasking without the fear of stumbling and falling (2007) and without the fatigue and discomfort of excessive energy demands as with other prosthetics (U.S. Department of Veteran Affairs, 2012).

Moreover, in the BiOM AK, the absence of positive power in the X2 knee is compensated for by the powered BiOM ankle which propels the amputee with each plantar flexion (iWalk, 2012). This makes for an effortless ambulation and a pronounced decrease in metabolic demands in above knee amputees (2012). Though published studies about the exact energy demands of the BiOM AK are not available yet, a study concerning the C-leg, the predecessor of the X2, performed by Dr. Kaufman reveals a decrease in metabolic demands with the use of microprocessor controlled knees (Kaufman et al., 2008). Dr. Kaufman explains that the energy efficiency of patients using the C-leg improved by 2.3% (2008). He also notes that the overall energy expenditure of the users of the C-leg increased by 6% due to increased activity levels as a result of the safe and comfortable walking ability provided by the knee (2008). He also adds that there was an improvement in gait quality which contributed to the energy efficiency, in addition to reduced stumbling and more confidence that led amputees to become more active (2008). It should be noted that the C-leg is not as advanced as the X2 and the X2 being much more sophisticated and efficient, especially when combined with the BiOM ankle, can lead to higher energy efficiency levels.

3.3.3.4. Practicality of Usage

The main factor that renders the X2 knee practical to use is its intelligent system that requires no conscious control, thus facilitating its use (Edwards, 2011).

Another practical feature of the X2 is that it has five programmable modes with the facility of switching between modes (Edwards, 2011). As mentioned before, tapping the ground three times can switch the user to another mode (2011). The switching can also be achieved by the provided remote control device (2011). This feature is particularly useful for very active amputees who usually switch to other legs before engaging in strenuous activities like running or biking.

The X2 has a long battery life, which can last 4-5 days (Edwards, 2011). It is advisable, however to charge it on a daily basis. The charger attaches to the leg magnetically and it indicates the charge status of the battery by various LED lights (2011).

The X2 can support weights up to 149 Kg and can accommodate patients with long limbs and also patients with short stature (Edwards, 2011). These categories of amputees are considered difficult to fit with prosthetics. However, the X2, owing it to its shorter than usual length can accommodate them (2011).

The X2 is splash resistant and can thus be used safely in cases of accidental splash or a sudden rainstorm (Progressive Orthotics and Prosthetics, 2012) . However, when coupled with the BiOM ankle, contact with water should be avoided (iWalk, 2012).

3.3.3.5. Cost

The X2 costs around $32000 (Tan, 2012). When coupled with the BiOM ankle, the BiOM AK costs around $80000 (Wynn, 2012).

3.4. Upper Extremity Prosthetics

Upper fringe prosthetics are artificial devices that substitute for a missing arm or parts of an arm and attempt at restoring the functions of a biological arm to various degrees to achieve the basic functions of an arm, mainly moving the arm to a desired position in order to hold or carry an object and perform desired operations (Wilson. 1963). Upper body prosthetic devices can be divided into four basic types: Cosmetic prostheses, body powered prostheses, externally powered prosthesis and hybrid prostheses (Kelly, 2011).

- Cosmetic prosthesis is used mainly for a cosmetic appearance and it is almost non-functional (Kelly, 2011).

- Body powered arms, on the other hand, are moderately affordable and popular prosthetic devices that are controlled by the relative motion of two body parts by means of a cable and a harness (Kelly, 2011).

- Externally powered arms are prosthetics that are powered by a battery switch system or by muscle signals like the myoelectric arm (Kelly, 2011).

- Hybrid arms are a combination of body powered and externally powered systems, where a cable is used for the elbow movement and batteries operate the terminal device (hand or hook) (Kelly, 2011).

In this report we will assess the performance of the moderately priced body powered arm and the high cost, externally powered, intuitively controlled and very advanced Luke arm. Our analysis also includes the recent and promising design, the artificial muscle operated arm.

3.4.1. Conventional or Body-Powered Prosthetic Arms

Body powered prosthetic arm designs vary according to the degree of amputation. In our analysis we tried to summarize these types under three broad headings. We can thus differentiate between below elbow, above elbow and shoulder disarticulation cases. However, all three kinds comprise the following basic elements: A socket, a suspension system (mainly a harness), a cable system for control and a terminal device such as a hook or a hand (Kelly, 2011). A wrist unit and an elbow lock can be added if needed.

©Computer Shopper, 2012 February                                                                                                                                                           Figure 20:traditional functional upper limb  prosthesis

with cable-controlled split hooks for hands.

- Below Elbow Prosthetic Arms:

In below elbow cases, a socket (or a split socket for very short stumps) connects the arm to the stump (Wilson, 1963). Its length depends on the degree of amputation (1963). The arm ends in an optional rotating wrist followed by the terminal device which could be a hook or a hand (1963). Though the hand is more cosmetic, the hook is more popular among amputees due to its high practicality and efficiency (1963). Both types of terminal devices operate by the same methods which can be either voluntary opening or voluntary closing (1963). The voluntary opening option is more popular and the hook is only opened to release an object, while it naturally stays in a nearly closed position when not in use (1963). The operation of the terminal device is performed by means of a cable connecting the hand or hook to a harness (preferably a figure 8 harness) (1963). A small tension, using the shoulder, the chest or other sound body parts involved in the design is enough to open the hook or the hand (1963). The below elbow amputee has the advantage of using his elbow and the remaining arm to move and position the arm to achieve the desired activity.

Figure 22:©http://www.oandplibrary.org/al/1963_02_001.asp                                                                                                                                                  Top: Below Elbow Prosthetic Arm and  Figure 8 Harness.                                                                                                                                       A, B, C: Loose Socket securing mechanisms

Figure 21:©http://www.oandplibrary.org/al/1963_02_001.asp                                                                                                                              Terminal Devices

- Above Elbow Prosthetic Arms

The above elbow prosthetic arm differs from the below elbow arm in a few details. The socket in this case is longer and covers the whole stump and it is connected most frequently to a figure 8 harness and in some cases to a chest strap (Wilson, 1963). Of major importance in above elbow cases is the elbow lock. In the absence of a

Figure 23: Above Elbow Prosthetic ©http://www.oandplibrary.org/al/1963_02_001.asp

biological elbow, an elbow lock helps the amputee to position his arm to achieve desired actions (1963). This lock is actuated by a cable connected to the harness (1963). A slight depressing of the shoulder is enough to lock the elbow (1963). The control of the terminal device can be achieved by the same cable or by adding another one (1963).

- Shoulder Disarticulation Prosthetics:

Shoulder amputees are the most problematic to fit and the corresponding device is usually heavy, excessive energy consuming and complicated to use to the degree that some patients switch to non-functional cosmetic types (Kelly, 2011). Shoulder disarticulation prosthetic devices are designed to individual needs and remaining body parts and their potential. Though a successful shoulder joint operated through the harness is a near impossibility, reasonably functional designs can be achieved using three cables and manipulating the harness (Mc Kenzie, 1965). In the case of bilateral amputees, further manipulations are needed (1965).

©http://www.oandplibrary.org/al/1963_02_001.asp                                                                                                                                                   Left:Bilateral Shoulder Disarticulation Prostheses and Special Harness         Right: Typical Shoulder Disarticulation Prosthesis

The following are the results of our assessment of the body powered prosthetic arm.

3.4.1.1. Elbow, Wrist and Fingers: Range of Motion, Articulation and Associated Scope of Functionality:

Body powered arm prosthetics provide a reasonable degree of functionality to amputees. The elbow range of flexion can reach around 135 degrees which is adequate for comfortable operation (Wilson, 1963). However, the functioning of the hook or the hand is far from the operation of a normal hand. The hook, as mentioned grips objects when closed and releases them when opened (1963). The hand is bulkier but it functions in the same manner with the movement of the thumb against two fingers (Wilson, 1963, Kelly, 2011). The rest of the fingers are usually non-functional (2011). However, with proper training, amputees can achieve multiple activities with these devices (Mc Kenzie, 1965). The hook, being simpler, provides for more visibility which facilitates operation Wilson, 1963). The amputee can orient the hook device as desired and lock it to facilitate operation, in case a wrist is absent (1963). In the presence of a wrist unit, a passive rotation and locking of the wrist help the amputee in manual operations (1963).

Moreover, body powered prosthetic arms are known for providing efficient sensory feedback through the socket-stump system, facilitating the operation of the terminal device Mc Kenzie, 1965). It is possible to hold delicate objects with hooks without crushing them, owing it to the sensory feedback mentioned (1965) and the high degree of visibility (Wilson, 1963). In addition, several task-specific terminal devices exist, such as ones for shaving, eating, golfing, fishing, hunting, skiing, playing musical instruments like the violin and even for driving cars (Kelly, 1994-2012, Mc Kenzie, 1965).

In order to achieve desired activities, an amputee should be able to move the elbow and the whole arm to the desired position and location (Wilson, 1963). This is easy for below elbow amputees who use the biological elbow and the residual arm muscles (Mc Kenzie, 1965). However, above elbow amputees can face difficulties. Body powered prosthetics are provided with elbow locks to facilitate operation (1963). The lock, as all other operations are activated by the cable harness and body movement complex (1963). This imposes restrictions on above elbow amputees who for example, can’t move the prosthetic arm to behind the head (Mc Kenzie, 1965). The harnessing impedes such large ranges of activity (1965). The picture is even grimmer for shoulder amputees who need to articulate three joints with a complicated cabling system and restricting harness to achieve movement (1965). In addition, body powered prosthetic arms can only provide the operation of one joint at a time. Thus, moving the whole arm takes several steps involving the moving of the wrist, elbow, shoulder and eventually the terminal device (Kuiken et al., 2009).

However, as Dr. Mc Kenzie explains, with proper harnessing and the hook as terminal device, body powered prosthetic arms with their sensory feedback, are reliable devices which provide independence and aid the amputee to lead an active life (Mc Kenzie, 1965). According to him, this applies even to bilateral amputees, who are able to do a variety of operations with their hook arms, with the exception of a few activities if given proper training (1965). In his paper, The Clinical Application of Externally Powered Artificial Arms, he states that even shoulder amputees who have to deal with the restrictions of a complicated and energy demanding harness-cable system, can achieve a certain degree of independence and tend to their basic needs (1965).

3.4.1.2. Control and Powering System

Control the terminal device and/or the arm in body powered prosthetic arms is achieved through one or more cables connecting different parts of the arm to the harness (Wilson, 1963). In below elbow cases, the cable is hooked to the terminal device (1963). In above elbow cases, an additional cable moves the elbow, though one cable achieves the movement of both the terminal device and the elbow, in alternative designs (Wilson, 1963, Mc Kenzie, 1965). For shoulder amputees, each of the three joints’ movement is achieved through the relative movement of body parts, such as shoulder extension or depression, chest extension and similar moves depending on the amputees muscle potentials, the design and the corresponding connection to the harness (Wilson, 1963, Mc Kenzie, 1965).

The body powered prosthetic arms as the name implies, are powered by the body muscles. The energy required to move these arms is provided by the amputee’s muscles (Wilson, 1963). This energy is of reasonable amount for below and above elbow amputees, but it becomes excessive for shoulder disarticulation cases where the absence of the shoulder joint, in addition to the elbow, the wrist and the hand, requires the movement of three joints using body movements which are further restricted by the heavy and complicated harness (Kelly, 1994-2012, Mc Kenzie, 1965).

3.4.1.3. Weight and Comfort

Body powered prosthetic arms are moderately light in weight (Kelly, 1994-2012), the more comfortable designs being the below elbow devices which can even make use of loose fitting sockets (Wilson, 1963). The main problem faced by amputees using the body powered arms is the harness and the associated discomfort. While some amputations require less harnessing, most above elbow cases require the figure 8 and even a harness with a chest strap to achieve demanding activities (Wilson, 1963). The latter are particularly problematic for women as they interfere with the breasts (Kelly, 1994-2012).

Discomfort faced by shoulder amputees is even greater due to the increased weight of the whole system, the restricting harnessing and the cognitive demand to operate the system (Kelly, 1994-2012). Moreover, these amputees have fewer sites for harnessing due to various missing body parts, depending on the case (Mc Kenzie, 1965). Shoulder amputees experience great discomfort and frustration and the use of their prosthetic arms tends to be exhausting and provides limited movements (1965).

3.4.1.4. Practicality of Usage

Body powered prosthetic arms need basic training before their adaptation to daily activities Mc Kenzie, 1965). Training may take several months and it needs to be done under the surveillance of a professional (Wilson, 1963).  The training is cognitively demanding and the time needed for training may stretch for longer periods in case of complicated designs (Mc Kenzie, 1965). The whole process proves frustrating for some above elbow and shoulder unilateral amputees who switch to a non-functional cosmetic arm to fill the sleeve and depend on one arm for operation (1965). However, once the device use is mastered, amputees can lead an independent life and even pursue hobbies, provided that he has enough muscle force for achieving arm movement and grip (1965).

Nevertheless, some patients find body powered prosthetic arms inefficient compared to the effort used to produce arm functioning. They also complain about sweating inside the socket which makes it slippery and insecure (Adee, 2008).

Another factor that reduces the practicality of these devices is the far from cosmetic and normal appearance due to the harness and cables (Harrison, 2008).

Despite the fact that they are not easy to use and that they have a far from acceptable appearance, body powered arms are popular, mainly because they are durable and require minimum maintenance (Kelly, 1994-2012). Even terminal devices exist in plastic covered varieties, making them very practical to wash (Wilson, 1963).

Moreover, the variety in the task-specific terminal devices allows amputees to achieve a variety of desired activities (Kelly, 1994-2012).

3.4.1.5. Cost

Body powered prosthetic arms vary greatly in cost. This is due to the fact that they are designed and fitted according to individual needs. However, the literature reviewed suggests that they can cost around $10000 (Mc Gimbsy, & Bradford, n.d.) which is considered a moderate cost if compared to other prosthetic arms (Kelly, 1994-2012).

3.4.2. The Myoelectric Arm Coupled with Targeted Muscle Reinnervation: The Luke Arm

Before beginning our assessment of the revolutionary Luke arm, a general overview of the myoelectric arm and targeted muscle reinnervation are necessary for a thorough understanding of the different features of the Luke arm discussed in the review.

3.4.2.1. The Myoelectric Arm

The myoelectric arm is an externally powered prosthetic arm that contains sensors, processors and motors (Dailami, 2002). It is controlled by electric signals called electromyograms or EMG resulting from the voluntary contraction or twitching of residual muscles (Dailami, 2002, Clement, 1998-2012). These signals are, in turn, recorded by adjacent sensors in the arm and transmitted to a controller that processes them and turns on the motors, thus sending the arm into desired motion (Clement, 1998-2012). However, the amputee is able to move one joint at a time and the simultaneous movement of several joints is not possible with the ordinary myoelectric arms (Otto Bock, 2012b). In clearer terms, the amputee needs to twitch several muscles, usually antagonistic, and in sequential order, in order to achieve a movement (2012b).

The myoelectric arm, contrary to the body powered arm, is powered by a battery and moved by motors (Dailami, 2002). Thus, it does not require the amputee’s muscle powers for operation (2002). A simple twitching of a residual muscle is enough to start the arm (2002).

However, the efficiency of the myoelectric arm depends on the quality of signals generated by the amputee’s muscles (Dailami, 2002). These signals should be interference-free and strong enough to be recorded and transmitted (2002). Consequently, it is suitable only for amputees with strong and sufficient residual muscles.

3.4.2.2. Targeted Muscle Reinnervation

Figure 25: TMR                                                                                                                                                                                           ©Rehabilitation Institute of Chicago

Targeted muscle reinnervation (TMR) is a surgical intervention devised by Dr Todd Kuiken, physician and engineer at the Rehabilitation Institute of Chicago, in order to improve the control and operation of myoelectric arms (Kuiken et al., 2009). As Dr Kuiken explains in his study Targeted Muscle Reinnervation for Real-time Myoelectric Control of Multifunction Artificial Arms, “The goal of targeted muscle reinnervation is to create a new surface electromyogram (EMG) signals that can control a motorized arm” (2009). TMR interventions involve the delicate surgical procedure of transferring the remaining nerves in the amputee’s residual arm to the chest, upper arm or shoulder muscles where these rerouted nerves are allowed to grow and branch (2009). Consequently, all brain signals that once controlled the arm will be sent to these new locations, since the brain retains the perception of the arm even if the arm is amputated (Otto Bock, 2012b). The muscles in the new location will contract with every brain signal ordering the arm to move and thus, additional or amplified myoelectric signals (EMG) are obtained (Kuiken, 2009).  The amplified EMG generated is then recorded by electrodes placed on the skin corresponding to the new nerve sites (2009). These signals are in turn transmitted into the prosthetic arm for functioning. This whole process allows the amputee to operate a myoelectric arm intuitively and the amplified brain signals enable him to move several joints at a time, just like a biological arm (2009). Therefore, TMR renders the operation of the myoelectric arm intuitive, simpler, faster, smoother and more efficient in that more than one joint will operate with one amplified signal and without effort (2009), so that the movement and operation of the arm are similar to that of a biological arm.

TMR establishes a neural-machine interface by making use of the body’s natural signals and as Dr. Douglas Smith, the first surgeon to perform TMR outside Chicago, explains, once performed, there is nothing that can ruin it as it is the case with chips inserted in the body and the brain and that can break or have to be replaced (Fairley M., 2007).

TMR is especially useful for bilateral shoulder disarticulation and for above elbow amputations, where generating strong and interference-free EMG for the operation of a myoelectric arm is highly improbable due to weak or absent muscles (Kuiken et al., 2005).

3.4.2.3. The Luke Arm

In the past few decades, many bionic arms have been designed by various research centers such as the i-Limb, the Smarthand, the Luke arm and many others (Kroeker, 2011). Our analysis of bionic arms which are advanced forms of myoelectric and/or externally powered arms will focus on the Luke arm, because it has been successfully coupled with TMR and has been tested extensively.

Figure 26: The Luke Arm                                                                                                 ©http://www.dekaresearch.com/images/deka_arm.jpg

The Luke arm was designed by Dean Kamen and a group of DEKA engineers. It was funded extensively by DARPA with the aim of providing a better life for U.S. war veterans and soldiers who had lost their arms in combat (Adee, 2008).

The Luke arm comprises a highly advanced technology that involves new control systems, as well as an advanced neural interfacing system (Kroeker, 2011). It takes its name from Luke Skywalker of the movie Star Wars who wore an advanced arm after losing his arm in a war (Adee, 2008).

With its many microprocessors, innovative sensory system, advanced socket and high degree of articulation (CBS Interactive Staff, 2009), the Luke arm has proven to be an innovation that can change the life of amputees. It was designed so as not to require any invasive surgery for operation (Adee, 2008). However, with Dr Kuiken’s revolutionary TMR procedure, the Luke arm was transferred to a whole new dimension, namely, purely brain controlled prosthetics (2008).

The following are the results of our assessment of the Luke arm.

3.4.2.3.1. Elbow, Wrist and Fingers: Range of Motion, Articulation and Scope of Functionality:

Unlike any other prosthetic arm, the Luke arm of DEKA provides for 18 degrees of freedom which is very close to the 22 degrees provided by the human arm (Adee, 2008). The shoulder, elbow, wrist and the fingers, with all their joints, are capable of movement in various directions and planes (2008). The arm, as a whole provides for a large variety of activities previously impossible for amputees. For example, the powered shoulder enables the shoulder amputee to reach over his head or perform lateral raises (2008). The highly articulated fingers and wrist enable him to perform delicate tasks like inserting a key in a keyhole and opening a door and even picking an object as small as a grape and eating it (2008).

Each part in the arm contains separate electronics and each part is powered by motors to achieve maximum functionality (Adee, 2008).

Moreover, the thumb is provided with a sensor that provides sensory feedback to the amputee through a tactor (vibrating motor) worn against the skin, anywhere in the body, such as the waist (Adee, 2008). As the grip on an object increases, the frequency of vibrations in the tactor increases informing the amputee of the amount of pressure needed for the grip in order not to damage the held object (2008). Consequently, the amputee is able to hold a plastic or paper cup without crushing it and even peel a banana without squishing

it (2008).

Figure 27: The Luke Arm                                                © www.engadget.com/media/2008/02/kamen-luke-arm

The features of functionality so far can be achieved by the Luke arm without TMR, with the amputee controlling the device and moving each joint to achieve a complete operation with muscle moves and foot pads (Adee, 2008). However, when coupled with TMR, the whole operation is achieved in a single smooth movement, as the shoulder, elbow, wrist and fingers move in coordination and simultaneously (Kuiken, 2005).

In short, the Luke arm, in its high degree of articulation, advanced motor control and sensory feedback, provides for a high degree of functionality.

3.4.2.3.2. Control and Powering System:

The Luke arm is designed for any type of control. However, the control used by DEKA is achieved through foot pads inside the shoes and shoulder pads connected to the arm by cables and the amputee is trained to achieve efficient control of the arm using the pads (Adee, 2008). Though, this control mechanism might seem unpractical, amputees who are given the chance to try the Luke arm express their satisfaction with it (Gould, 2012). Nevertheless, as one amputee explains, the Luke arm is best used when standing and not while walking due to the pads in the shoes (2012). However, when coupled with TMR, the Luke arm controls become completely natural (Kuiken, 2009). The control is then achieved intuitively, as explained in the TMR section.

The efficiency of the intuitive control of the Luke arm was studied extensively by Dr. Kuiken and his colleagues. In one experiment where the speed of initiation and completion of motion performed by amputees who had undergone TMR surgery, including five shoulder disarticulation amputees and five transhumeral amputees, given 10 different tasks to be achieved using a virtual arm with the same specifications as an advanced prosthetic arm, Dr Kuiken concludes that TMR produces EMG signals for the real time control of advanced prosthetic arms (Kuiken, 2009). The results of his experiment are summarized in the table below:

|  | TMR Patients | Control(non-amputees) |
| --- | --- | --- |
| Motion Selection Time Elbow and Wrist in seconds | 0.22 | 0.16 |
| Motion Selection Time Hand Grasp in seconds | 0.38 | 0.17 |
| Motion Completion Time Elbow and Wrist in seconds | 1.29 | 1.08 |
| Motion Completion Time Hand Grasp in seconds | 1.54 | 1.26 |
| Motion Completion Rate % (for 108 trials) | 96.3 | 100 |
| Hand Grasp % (for 72 trials) | 86.9 | 96.7 |

Table 5: TMR                                                                                                                                                                                                                  Adapted from Targeted Muscle Reinnervation for real Time Myoelectric Control of Multi- Function Artificial Arms by Kuiken et al., 2009

The same experiment includes the further use of the Luke arm by three amputees who had undergone TMR. The results, as Dr Kuiken concludes are highly encouraging and TMR enables the fast and smooth control of the Luke arm (Kuiken, 2009).

In another experiment performed by Dr Todd Kuiken and Dr Gregory Dumanian, the surgeon who performed the TMR operations, an amputee with bilateral shoulder disarticulation and another with transhumeral amputation, were tested for speed and efficiency in controlling a myoelectric arm after TMR surgery (Kuiken, 2005). The bilateral patient was 26% faster on his TMR side and the transhumeral patient was 84% faster (2005). The performance of both amputees was also improved with TMR (2005). This is again due to the amplified EMG signals that enable the amputee to move several joints in the arm simultaneously with their brain signals (2005).

These studies by Dr Kuiken and his colleagues suggest that the Luke arm performs best with TMR and brain signals that make the control simple and effortless and achieve smooth and real time motion. However, in the absence of TMR, the Luke arm mechanical control comprising food pads and muscles, is efficient (Adee, 2008), especially that it saves amputees from the cognitive efforts required to manipulate muscle contractions in ordinary myoelectric arms (Otto Bock, 2012).

Finally, it should be mentioned that the power needed for operating and moving the Luke arm is supplied by a Lithium rechargeable battery inside the arm and no muscle energy is needed to operate it (Adee, 2008).

3.4.2.3.3. Weight and Comfort Factors

The Luke arm despite its electronic circuits, battery, motors and other components weighs only 3.6Kg which is the weight of the arm of an average size woman’s arm (Adee, 2008). Being lightweight amputees find it comfortable to use.

The comfort factor is further enhanced by the innovative socket system which includes tiny bubbles or balloons that inflate when the arm is in use and thus tighten the socket (CBS Interactive Staff, 2009). When relaxed, the socket loosens automatically (2009).  Moreover, because of this secure socket system, it is enough to strap the arm on across the chest and use it (Adee, 2008). This innovation saves amputees the discomfort of harnesses and the sweating inside sockets that make the arm slippery and require very tight sockets (CBS, Interactive Staff, 2009).

3.4.2.3.4. Practicality of Usage

The Luke arm, when strapped on, is highly practical to use as confirmed by several amputees involved in the DEKA project and participating in testing the arm for further improvement. The controls of the arm are simple and can be learned in around 5-10 hours of practice (Adee, 2008). Moreover, the sensory feedback makes it even more practical (2008).

When coupled with TMR, the use of the arm becomes even more practical, but it still needs training to achieve efficient results (Fairley, 2007). The only inconvenience associated with TMR, is the time required for nerve reintegration after surgery, as it is around 0.5-1.5 years, though some patients required only six months (Otto Bock, 2012b, Kuiken, 2005).

Another feature that makes the Luke arm practical, is that it is modular, meaning each part is separate from the other (Adee, 2008). Consequently, it can be customized to individual needs and the amputee can wear only the parts necessary for his amputation and not the whole device from hand to shoulder (2008).

The final product is covered with synthetic material colored to match the skin color, thus giving the arm a natural and cosmetic look.

One inconvenience, though not mentioned in the literature, but evident in the videos is the noise associated with the movement of the arm, but it can be considered minor when compared to what the arm can accomplish especially with TMR. The Luke arm as described by the test pilots and Fred Downs, the VA official in charge of prosthetics, feels like a natural extension of the body (CBS Interactive Staff, 2009).

3.4.2.3.5. Cost

The Luke arm has not been marketed yet, but as stated by DEKA representatives, it should cost around $100000 and probably more (Saenz, 2009), which is not surprising, since most advanced arms are in the same range. However, this cost does not include the TMR surgery which is a complicated and delicate procedure performed only in a few hospitals. TMR raises the cost of the Luke arm by around $300000(Ryerson University, 2011).

3.4.3. The Artificial Muscle Operated Arm (AMO)

The Artificial Muscle Powered Arm or AMO is a multi-award winner innovative prosthetic device designed by two biomedical engineering students, Michal Prywata and Thiago Caires, at Ryerson University, Canada, whose aim was to design an arm that was as easy to purchase as a cell phone (CBC, 2011). It is a low cost brain controlled arm powered by artificial muscles and compressed air (2011). It is also simple, has no electronic circuits and does not require any prior surgery to operate (Ryerson, 2011).

Figure 28 AMO: Headset and Arm                                                                                                                                                                    ©Ryerson University

The AMO is a two piece device that comprises the arm and a headset (Ryerson University, 2011). The head-piece is worn on the head and contains multiple sensors that receive, record, and amplify brain signals (Hein, 2011). The arm includes eight artificial muscles in the arm-piece, one in the shoulder, one in the elbow, one in the wrist, one in the thumb and four in the fingers (2011). It also contains a microprocessor that receives the signals coming from the head-piece wirelessly and processes them (Ryerson, 2011). In addition, it has a pneumatic pump connected to a tank of compressed air (2011). With the simple components mentioned above, the AMO arm operates when brain signals are transmitted via the sensors to the processor which, in turn matches them to available database and sends out a new signal to operate the pneumatic system (2011). Accordingly, certain valves open and fill corresponding muscles with air leading to their contraction and resulting in the movement of the arm (Discovery Channel, 2011).

The AMO arm hasn’t been clinically tested yet and our assessment of its different features is based on media reviews and videos, in addition to a short interview with Dr. N. Sabah, professor of biomedical engineering at the American University of Beirut. The following are the results of our assessment of the AMO arm.

3.4.3.1. Elbow, Wrist and Fingers: range of Motion, Articulation and Scope of Functionality

The AMO arm in its initial design can perform four basic functions (CBC News, 2011):

- It can raise and lower the arm at the shoulder level.

- It can flex the elbow.

- It can rotate the wrist.

- It can close the fingers to grab an object.

However, individual finger motion is not possible and the four fingers move in one move, though the designers are currently working on the issue (Ryerson University, 2011). They have also promised to provide the fingers with sensors to provide sensory feedback and allow the user to control the grip force for gentle handling

of objects (2011).

Figure 29: The AMO Arm                                                                                                 © Ryerson University

In its current condition, the AMO arm can provide for basic grabbing and carrying objects and the artificial muscles can withstand up to 200 pounds of loads (Discovery Channel, 2011). However, the AMO arm can’t achieve picking up a key, inserting it in a lock and unlocking a door, due to lack of individual finger articulation control, but as mentioned, these features will be available in the near future (Ryerson University, 2011).

3.4.3.2. Control and Powering System

What differentiates the AMO arm from existing prosthetic arms is its control and powering system. It is controlled by the brain (Ryerson University, 2011). The brain as mentioned in our analysis of the Luke arm retains its perception of the arm even if the arm is inexistent. Consequently, it sends out signals or biopotentials (voltage differences) ordering the arm to move (Discovery Channel, 2011). These signals are transmitted by the sensors in the headpiece to the microprocessor which in turn compares them to stored database and sends out the appropriate match signal which operates the arm in the desired direction and movement (2011). It should be mentioned that the database comprises recorded signals of normal arm functioning and it took the designers one whole year of programming work to achieve it (Ryerson University, 2011). As explained in the introduction, the pneumatic pump is initiated into action by the new signal coming from the database and the required muscles fill with air coming from the tank and the arm is set into motion (Hein, 2011). This whole process indicates that the AMO arm is purely brain controlled.

However, according to Dr. N.Sabah, the superficially located electrodes may not be able to record all brain signals for desired arm movement, but training helps to overcome this problem. The arm, on the other hand, as Dr. Sabah explains, can be subject to electromagnetic interference that can affect the wireless connection (Sabah N., Personal Communication, December 18, 2012).

As for the powering of the arm, the AMO arm has no motors, and does not require electric or battery power to operate (Ryerson, 2011). It is powered solely by the compressed air it utilizes (2011). The microprocessor, on the other hand, as Dr. Sabah clarifies, needs a power source, which could be a battery or even power generated from the compressed air itself, but the few media reviews about the AMO arm do not specify how the microprocessor is powered (Sabah N., Personal Communication, December 18, 2012).

3.4.3.3. Weight and Comfort factors

As mentioned, the AMO arm being a recent innovation, it hasn’t been put to clinical testing yet and the media reviews don’t refer to its exact weight and the associated comfort factors. However, in the absence of motors and electronics, the arm should be lightweight if compared to other myoelectric arms. Moreover, the arm doesn’t require any restricting harnesses and it can be used by strapping it on across the chest (Hayes, 2011). This feature is a relief for all amputees who complain of the discomfort caused by traditional harnesses.

3.4.3.4. Practicality of Usage

Like all brain controlled prosthetics, the AMO arm requires training for efficient operation. The basic functions of contracting, extending the arm, and closing the hand in a grip can be achieved in less than ten minutes of practice as demonstrated by a Discovery Channel reporter (Discovery Channel, 2011). However, an additional few weeks are needed to establish flawless control of the brain-arm link in order to perform complex tasks (Ryerson University, 2011).

The wireless connection between the head-piece and the arm can be considered a feature that makes the arm practical to use, though wearing a head-piece can be prove to be annoying or not cosmetic for some patients. However, as clarified by Dr. N. Sabah, the headset should not cause any discomfort, since it is noiseless and signals picked from other sources are negligible. Moreover, according to Dr. Sabah, the electrodes are safe to use. They do not affect the brain in any way and do not cause any brain damage or related ailments (Sabah N., Personal Communication, December 18, 2012).

The arm could also prove useful for wheelchair patients who can use it to reach out and grab objects without moving their bodies (Ryerson University, 2011).

The power source which is the air tank lasts 16 hours and future designs will include the air tank in the arm for more practicality (Discovery Channel, 2011).

3.4.3.5. Cost

With no electronic circuits and expensive parts, the AMO arm costs around one fourth the price of other brain controlled prosthetic arms (Ryerson University, 2011). Moreover, it saves patients the high cost of TMR surgery, which is still not available except in a few hospitals around the world (Ryerson University, 2011). The arm should be ready for marketing in 2013 and the market price of the AMO arm, as stated by the designers, will be around $25000, which makes it an affordable prosthetic arm, as intended by its designers (Hayes, 2011).

- Discussion of Results

  - Below Knee Prosthetic limbs

As mentioned in our general overview of prosthetic feet, current below knee prosthetics can be divided into conventional and advanced prosthetics. Our results include the assessment of the conventional SACH foot and Jaipur foot, in addition to the advanced BiOM ankle. The following is a detailed discussion of the advantages and disadvantages associated with the use of these prosthetic devices.

4.1.1. The SACH Foot

The SACH foot, as shown by our results, is a widely used conventional prosthetic foot. Its use is accompanied with several advantages, especially for low activity and weak amputees, in addition to a wide range of disadvantages, particularly for active amputees.

4.1.1.1. Advantages

The SACH foot, with its solid ankle and cushioned heel owes its popularity mainly to its low cost (Nobbe Orthopedics, 2001) and to the solid structure of the foot that acts as a stable platform for balance (Carpenter, Hunter, & Rheaume, 2008). The compressible heel that comes in three stiffness variations accommodates different needs and acts as a shock absorbent, in addition to contributing, to a small degree, to the loading of the foot for achieving toe off, thus facilitating the movement of the leg (Gordon, & Ardizzone, 1960). Moreover, the lack of articulation makes the foot rigid and stable and thus, suitable for the elderly and the weak, as well as amputees in rehabilitation stage (Marinakis, 2004). Hence, it is especially suitable for amputees with low activity levels and it gives the user stability and confidence.

Another advantageous feature of the SACH foot is the practicality of its usage. It is a lightweight (Ottobock, 2012a) and durable foot (Adalarasu, 2011), and has a simple design that has no movable parts and does not require regular maintenance, and thus has a very low maintenance cost (Hill, n.d.). This feature makes it especially useful for amputees residing in remote villages and who have difficult access to prosthesis services and clinics.

Moreover, the SACH foot is durable with a life span of 2.5 years and can withstand humidity to a certain degree (Ottobock, 2012a). It is lightweight and exists in a variety of sizes to fit children, women and men and all in natural looking designs (2012). Bearing in mind that the SACH foot, like most conventional prosthetics, requires amputee muscle work to move the leg, the lightweight is a comfort factor that eases the lives of amputees and reduces muscle and joint strains imposed by prosthetics.

One final advantage associated with the SACH foot is the cost factor. The SACH foot is a low cost foot (Nobbe Orthopedics, 2001) and the fitting can be done for a few hundred dollars as it can reach $8000 (Jaipur Foot, 2007), depending on the make and quality of parts. Thus it can accommodate different economic levels of amputees.

4.1.1.2. Disadvantages:

Unfortunately, the low cost and low maintenance SACH foot is not without disadvantages. As shown by our results, the solid ankle of the SACH foot provides for no articulation (Gailey, 2005). While this rigidity is desired for stability and is suitable for low activity users (Carpenter, Hunter, & Rheaume, 2008), it pauses major problems for active amputees by limiting their range of activities and the speed of ambulation. For example, the lack of real dorsiflexion results in a weak loading response, making walking tiring for amputees (Gailey, 2005). Moreover, the rigidity proves to be impractical for driving (pushing the accelerator with the prosthetic foot), climbing and descending stairs and managing ramps, as these activities prove to be tiring and demand continuous and excessive muscle work (Gordon, & Ardizzone, 1960). In addition, the rigidity of the foot and the lack of real dorsiflexion, result in low self selected or preferred speeds, making the SACH foot not suitable for high ambulation speeds (Gailey, 2005). The lack of real dorsiflexion also results in a delay in loading on the prosthetic side (Edelstein, 1988). As explained in our results, the time spent by the prosthetic foot between heel strike and toe off is twice that spent by the normal side (1988). This delay not only results in slower walking speeds, but also in an unnatural gait that becomes highly observable at higher speeds, where excessive hip and knee use compensate for the shortcomings of the foot (1988). Thus, we can conclude that the SACH foot, due to its rigidity is not suitable for active amputees.

Moreover, the lack of real lateral articulation can’t be compensated for by the extremely limited degree of lateral movement provided by the cushion heel that extends to the sides (Gordon, & Ardizzone, 1960). Thus, the SACH foot can’t provide for a safe and easy walk on uneven or rough surfaces and is better confined to level ground (1960). This limitation makes the SACH foot inconvenient for outdoors use and hence, inconvenient for active amputees who wish to engage in an active and normal community life.

Besides the limitations in versatility of mobility and walking speeds, the use of the SACH foot proves to be inefficient in terms of metabolic demand. The SACH foot, as shown by our results, provides for a maximum of 30% energy return efficiency, which is very low and can’t provide for an effortless or easy ambulation (Barr et al., 1992). This efficiency further decreases with increased speed obliging the amputee to produce more work using his hip extensors and leg muscles and in the process leads to increased use of the hip and hip displacement to push the amputee forward (Edelstein, 1988, Gailey, 2005)). In simple terms, the metabolic demand associated with the SACH foot is high and consequently, its use causes fatigue, making long hours of activity difficult and tiring for the amputee who literally has to carry his foot and push it forward all through his ambulation.

The use of the SACH foot is also associated with some minor limitations which make it inconvenient for some amputees. One such limitation is the 125Kg weight limit (Ottobock a, 2012), bearing in mind that overweight is a common problem for amputees who usually lead inactive lives, especially when using conventional prosthetics. Another disadvantage, though less significant, is the fact that the SACH foot can’t be used barefoot because it comes in specific heel heights and it should be used with the corresponding shoes, thus limiting its practicality of usage (Ottobock a, 2012). In addition, the practicality factor is further aggravated by the time needed for fitting, which can stretch up to three months requiring multiple visits to the prosthetist (Jaipur foot, 2007).

4.1.1.3. Conclusion

The SACH foot can serve as prosthesis of choice for low activity amputees weighing below 125Kg and it is ideal for weak amputees who benefit from the high level of stability provided by the foot and the accompanying confidence. It is also a choice to be considered when low cost and low maintenance are deciding factors. However, the SACH foot is not advisable when high walking speeds, outdoors use, high range of activities and gait quality and symmetry are desired, making it not suitable for high activity amputees who wish to live a productive and normal life. In short, the SACH foot replaces the amputated foot, but does not restore a normal life for amputees.

4.1.2. The Jaipur Foot

The Jaipur foot, as shown by the results of our review of the literature, the Jaipur organization site, user comments and videos, is a low cost prosthetic foot that is considered the prosthesis of choice for developing countries (Prahald, 2003). Its use is accompanied by numerous advantages, especially when compared to its very low cost, making it ideal for low income amputees and NGOs.

4.1.2.1. Advantages

The Jaipur foot costs only $5 and the whole prosthesis comes at a final price of $30-35 (Jaipur Foot, 2007). It owes its low cost to its cheap but durable material, as it is made of cast away tire rubber and polyethylene pipes (Prahald, 2003). Nevertheless, being low cost doesn’t make the Jaipur foot less efficient than higher cost prosthetics. On the contrary, our results show that with the various advantages it offers, the Jaipur foot matches high cost flexible prosthetics such as the $3700 Variflex (2003).

The Jaipur foot owes its numerous advantages to its high degree of articulation in many planes (Mathur). Its capacity for 40° of dorsiflexion results in a reasonable degree of loading response and provides for an efficient toe off and an easy walk (Prahald, 2003). It also enables users to engage in a wide variety of activities such as driving, biking, squatting, crouching, climbing stairs and trees and even jumping (Prahald, 2003, Jaipur Foot, 2007). Consequently, the Jaipur foot, unlike the SACH foot, is suitable for highly active amputees.

The lateral and transverse articulations of the foot that provide for 10° of eversion and inversion, 7° of supination and 5° of pronation, provide for stability on uneven, rough and muddy terrains by enhancing equilibrium (Prahald, 2003). The transverse articulation also enables cross-legged sitting, which is not possible with rigid prosthetics like the SACH foot (Mathur, n.d.). Moreover, the balance and stability provided by the various movements of the foot in different planes is enhanced by the large surface area of the foot that acts as a stable platform for bearing the body weight in all stages of the gait cycle (Kapp, & Cummings). These features make the Jaipur foot ease the life of any below knee amputee and enable him to lead an active life.

Besides a high degree of versatility of mobility, the Jaipur foot, unlike the SACH foot, accommodates for variable walking speeds, enabling fast walking and running when needed (Mathur, n.d.). Our results show that the associated gait quality is stable and close to normal, owing it to the high degree of transverse articulation and the efficient dorsal flexion that absorb ground reaction forces and distribute them evenly (Jaipur Foot, 2007, Sethi, Udawat, & Kasliwal, 1978). Moreover, the total contact socket regularly used with the Jaipur foot aids in distributing these forces evenly over the stump, rather than localizing them, and thus enhances comfort and reduces pain, making walking, running and even long hours of standing comfortable for amputees (Mathur, n.d.).  The high articulation capacity of the Jaipur foot and the efficient pressure distribution provide also for metabolic efficiency, thus facilitating mobility and reducing fatigue (Mathur, n.d.). The metabolic efficiency is further enhanced by the spongy constitution of the foot which gives it a pneumatic joint attribute and provides the amputee with additional thrust and energy transmission across the leg (Sethi, Udawat, & Kasliwal, 1978). These features are of high importance for an active and productive life and the Jaipur foot, unlike the SACH foot which cannot accommodate to comfortable walking for long hours especially at increased speeds and cannot provide for metabolic efficiency, proves to be a good choice for amputees who need to work long hours and walk long distances.

The numerous advantages of the Jaipur foot also extend to its practicality of usage. The Jaipur foot is a durable prosthetic with a life span of three years (Adalarasu, 2011). It can withstand water, mud and all road and weather conditions (Jaipur Foot, 2007). It comes in three shades of skin color and has a cosmetic appearance (2007). It can be used barefoot, as well as with shoes (2007). This high degree of practicality is accompanied with its lightweight, which is a comfort factor. Our results indicate that at 850g of weight for the foot and 3.11Kg for the whole device, the Jaipur prosthetic is lighter than the average weight of the lower leg of a 55Kg person which is 3.36Kg (Jaipur Foot, 2007). This lightweight makes it easier for the amputee to move the leg and reduces the additional effort required by his body and leg muscles, thus enhancing comfort in ambulation and reducing fatigue. These features, though not all unique to the Jaipur foot, when combined to the previously discussed advantages make the Jaipur foot highly efficient in enabling amputees to lead a productive life.

Finally, a major feature praised by many users of the Jaipur foot, is the rapidity of making the prosthesis and fitting the amputee, which requires around 2-3 hours only Jaipur Foot, 2007). The Jaipur below leg prosthesis, unlike the SACH foot, is made and assembled on the spot and amputees require only minor training and are able to walk out of the fitting center wearing the Jaipur leg (2007). This fast fitting feature is rare with conventional prosthetics and it spares the amputee the multiple visits to the prosthetist and the accompanying cost, energy and time.

4.1.2.2. Disadvantages

The literature reviewed barely points to disadvantages associated with the Jaipur foot, the reason being the negligible cost of the foot accompanied with a high range of advantages. However, our observations of amputee gait in various short online videos show a slight limp and a slight hip displacement on the sound side, especially at increased speeds of walking. However, this is expected because the Jaipur foot despite its high degree of articulation doesn’t provide for plantar flexion which is essential for a normal and effortless ground clearance (Adalarasu, 2011).

In addition, as our review suggested, the Jaipur foot is still made on the spot and by people who haven’t been given enough training. Though highly advanced equipment and methods for alignment and socket fitting are used (Jaipur Foot, 2007), the absence of professional education and experience and the speed at which the prosthesis is made, doesn’t always come without problems, especially when considering the enormous numbers of amputees fitted daily at the BMVSS centers (Rosenberg, 2011). However, the BMVSS is currently working on a standard foot design and composition for mass production, in order to give the Jaipur foot uniformity (Sebastian, 2012). The organization is sparing no effort to enhance the quality of Jaipur prosthetics in every possible way (Jaipur Foot, 2007).

Another factor that can be considered a disadvantage concerning the Jaipur foot is that the Jaipur foot is not a commercial product and it is only available through the Jaipur Foot organization and NGOs (Jaipur Foot, 2007). Consequently, an amputee needs to travel to Jaipur or visit BMVSS camps in order to be fitted with the prosthesis. However, the recent collaboration of BMVSS and Michigan Technological University to take the Jaipur foot to mass production and establish it on the international level can solve this problem and make the foot available for all amputees (Sebastian, 2012).

4.1.2.3. Conclusion:

The Jaipur foot with its low cost and various advantages is prosthesis of choice for all below knee amputees. It enables amputees to lead an active life and restore their productivity, thus enhancing their quality of life. The high degrees of mobility, stability, speed accommodation, comfort and practicality, place the Jaipur foot ahead of the SACH foot, especially for active amputees.

4.1.3. The Power Foot: BiOM Ankle

As shown in the results section, the BiOM ankle is a robotic device that substitutes for a missing lower leg and replaces the action of the foot and the ankle, mainly the Achilles tendon and the calf muscles(iWalk, 2012). Through a highly advanced technology that mimics the biological ankle, it provides for a nearly normalized gait, making the BiOM ankle a natural extension of the amputee’s body (2012). In fact, as described by its inventor, Dr Hugh Herr, the BiOM ankle is a “technological embodiment” that transfers its user to an era where there is no more differentiation between what is natural and what is not (2012). The BiOM ankle, as an advanced technological device offers multiple advantages to amputees, but its use is also accompanied with certain disadvantages, though minor.

4.1.3.1. Advantages:

The BiOM ankle, with its high degree of articulation, advanced and smart control system of the position of the foot in swing phase, automatic adjustment of stiffness all through the gait cycle, in addition to the powered plantar flexion that propels the amputee forward with every step, provides for an efficient, consistent and normalized ambulation (iWalk, 2012, Au et al, 2007). It also provides for the same range of activities as those provided by a biological foot and ankle (2012). This versatility in mobility is also associated with a high degree of stability and balance enabling the user to walk safely on rough terrain, ascend and descend stairs with ease and manage ramps without effort and all in real time and with no delays between the sound leg and the prosthetic side, due to the highly advanced system of processors and sensors (iWalk, 2012, Prokaza, 2012). Thus, the BiOM ankle enables the amputee to lead an active life and pursue any desired activity.

The operation of the BiOM ankle is purely reflexive and does not demand any cognitive efforts (iWalk, 2012). The ankle simply thinks for itself and adjusts stiffness, position and power according to need, terrain and speed automatically (Au et al., 2007, Prokaza, 2012). Thus, the amputee doesn’t need to look down with every step and control his steps with his muscles as with other prosthetic feet (iWalk). Moreover, the fact that it responds automatically to all walking speeds gives the amputee a higher self-selected speed enabling him to keep up with normal ranges of walking and activity including running (Au et al., 2007). The power provided by the ankle is also reflexive and it increases with increased speed, leading to an effortless mobility (Eilenberg, Geyer, & Herr, 2010). These reflexive response features of the BiOM enable the amputee to engage in any desired activity and without anxiety. In fact, some amputees have reported forgetting their disability, and returning to their normal lifestyle (iWalk, 2012).

Another feature that highly benefits users is the highly symmetrical gait provided by the BiOM ankle (Au, & Herr, 2008). This symmetry is further enhanced by the personal tuning option that enables the adjustment of the ankle parameters according to personal needs via Bluetooth and provided software (Prokaza, 2012). The symmetry in gait provided by the BiOM ankle not only contributes to the amputee’s psychological well being and his self-confidence, but it also preserves the sound parts of the body such as the back, knee and hips which experience high degrees of strain with other prosthetics (iWalk, 2012).

The metabolic efficiency of the BiOM ankle is another feature that constitutes a major advantage for its user. As shown by our assessment results, the 30J of work produced with every step and the 67% energy efficiency of the ankle reduce the metabolic demands and even normalize energy expenditure by the amputee, especially at normal speeds (Herr, & Grabowski, 2011). Moreover, the force or power moving the foot is provided by the prosthesis and not by the amputee’s muscles, which further reduces metabolic costs, in addition to accompanying strain and fatigue (iWalk, 2012). Moreover, with its mere 2Kg of weight that is even cancelled by the propelling power provided by the motor at every plantar flexion, the BiOM further reduces the need for extra energy expenditure, reduces fatigue and provides comfort for its user (Herr, & Grabowski, 2011). The combination of powered steps, normalized metabolic demand, light weight and comfort results not only in the reduction of the usual fatigue associated with the use of prosthetic feet, but also reduces the pain experienced by amputees in different parts of the body, especially the knee, hip and back on the sound side (iWalk, 2012). In fact, as reported by several amputee soldiers, the BiOM ankle reduces the intake of pain killer medication to half (2012). Moreover, the resulting ease of activity decreases overweight and obesity rates common to amputees due to an inactive lifestyle (Knapp, 2012). Consequently, the BiOM ankle, in the long run, reduces overweight related ailments, such as heart and circulatory problems (Knapp, 2012). The health benefits are also accompanied by reduced health care expenses which is advantageous not only to the amputee, but also for government health care systems and health insurance companies.

In addition, as shown by our results, the BiOM ankle with its long battery life that can provide for almost a whole day of active use, the two spare batteries which at a weight of 0.22Kg can be carried in a pocket, and the intuitive interface comprising LED lights and a buzzing option to indicate charge status, the BiOM proves to be a practical device (Herr, & Grabowski, 2011, Fierce Medical Devices, 2012). The amputee can just switch an empty battery with a spare one and continue his normal activity.

Moreover, the short time needed to adapt to the ankle is also a positive factor that indicates the ease of use and efficiency of the foot, and relieves amputees from tiresome and frustrating training sessions. As shown in the results section, an amputee needs around 20 minutes to start walking comfortably with the BiOM ankle which can’t compare to the much longer training periods required by other prosthetic feet (Au et al., 2007).

4.1.3.2. Disadvantages:

The BiOM, like any other device is not perfect, but the disadvantages associated with this robotic ankle can be considered minor when compared to the whole range of advantages previously discussed. One disadvantage is that, being an electronic device, it can’t tolerate contact with water, dust and dirt, thus limiting its practicality of usage in daily life (iWalk, 2012). Another shortcoming is the weight limit. The BiOM can accommodate 170-250lb (2012), which can be an inconvenience for some amputees who have overweight problems resulting from inactivity imposed by the use of tiring conventional prosthetics. Moreover, the use of the BiOM ankle is associated with a bit of mechanical noise which can prove to be annoying or embarrassing for self conscious amputees (Strickland, 2012).

In addition, the cost of the BiOM ankle at an average of $50000 can make it out of reach for many amputees (Wynn, 2012). This inconvenience associated with the cost is further aggravated by the cost of required regular maintenance of parts. However, when compared to the normalized walking ability provided by the BiOM ankle and the associated psychological, socioeconomic and health benefits, the cost factor becomes less significant.

4.1.3.3. Conclusion:

In short, the multiple advantages of the BiOM ankle mentioned above indicate and prove its capacity at enhancing the amputee’s quality of life. The BiOM ankle has already enabled many U.S. below knee amputees to return to training and active duty, which is impossible to achieve with other prosthetics and particularly conventional ones (Keefe, 2011).

Finally, the BiOM ankle is a whole new dimension in the world of prosthetics. It is a robotic ankle that enables the amputee to regain control over his life and enhances his quality of life and most importantly, restores normal life.

  - Above Knee Prosthetic Limbs

As discussed in detail in our general overview of above knee prosthesis, the knee is the differentiating unit in above knee prosthetic limbs. Our assessment results include the affordable single axis constant friction knee, the polycentric Jaipur knee and the highly advanced BiOM AK with its robotic X2 knee. Our results reveal various advantages and disadvantages associated with each of the three mentioned prosthetics which will be discussed in detail in the following section.

    - The Single Axis Constant Friction Knee

Our results suggest that the single axis constant friction knee is the most basic and the most commonly used knee in developing countries due to its low cost (Dupes, 2008). It is a simple joint that operates like a door hinge and is controlled by a constant friction unit, with an optional manual lock that acts against knee buckling for safe standing (Dupes, 2008, Michael, n.d.). The single axis knee is a primitive prosthetic device and its use is accompanied with a few advantages and a long series of disadvantages and limitations. However, in developing countries like Lebanon, where amputees still use the $700 rigid knees, the single axis constant friction knee can prove useful.

4.2.1.1. Advantages

The single axis constant friction knee is a simple and affordable prosthetic device that does not require regular and costly maintenance and it is useful for people with difficult access to prosthetic centers such as those living in rural areas, as it is suitable for children who outgrow their prosthesis as they grow (Michal, n.d.). With its constant friction, the single axis constant friction knee can provide sufficient swing control and the optional manual lock can provide for the control of the leg when extended so as to prevent buckling when standing, walking on uneven ground or crossing obstacles (Michal, n.d.). The knee also provides for sitting with a flexed leg (Michal, n.d.).

Moreover, at a maximum cost of $500 for the knee and a maximum cost of $3000 for final fitting of the whole prosthetic leg, the single axis constant friction knee is considered low cost when compared to the cost of other above knee prosthetics that cost a minimum of $10000 (Orthotic and Prosthetic Appliances: Billing Codes and Reimbursement Rates-Prosthetics, 2012).

4.2.1.2. Disadvantages

Our results show that the main and perhaps the most significant limitation of the single axis constant friction knee is the fact that it depends excessively on the amputee’s muscles for control and since it has no stance control, its use involves continuous muscle effort to prevent knee buckling and to control the knee (Michal, n.d.). Thus it is only suitable for strong and able muscled amputees who are able to control the knee in walking, especially on uneven ground, while managing stairs and also while carrying loads. The knee even causes buckling problems on carpets and grassy land, where major effort is required to control its flexion (Michal, n.d.). However, many young and active amputees adapt to this knee, while the weaker and the elderly opt for the manual lock which locks the knee in extension and provides for a safe crossing of uneven terrain and obstacles, the result being a more stable and safe mobility with a very awkward rigid and tiring gait (Dupes, 2008). Therefore, in the single axis constant friction knee the lack of stance control is a major disadvantage and it greatly limits the range of activities and safe walking, especially for weak amputees, thus limiting mobility in general.

Moreover, the limited range of mobility is also associated with a constant speed of walking. The knee, being a constant friction knee, can be adjusted to only a single preferred speed of walking, thus offering no accommodation for varying cadence (Greene, 1983). The inability of the single axis constant friction knee in providing variable walking speeds is another disadvantage that stands in the way of restoring normal ambulation for amputees, thus limiting their quality of life.

The disadvantages accompanying the use of the single axis constant friction knee also extend to the associated gait quality, which as emphasized in our results, is a rigid and asymmetric gait exhibiting pronounced effects on the pelvis and other body parts (Farber, & Jacobson, 1995). The abnormality of the gait is further accentuated by the fact that the prosthetic limb has to be shorter in length than the sound limb, since the knee extends during walking and acquires a length exceeding that of the normal side (1995). As shown by our research results, the speed of walking associated with the single axis constant friction knee is only 2/3 of normal walking speed, while the accompanying abnormal gait pattern is characterized by wide and uneven strides on the prosthetic side with long stance periods and a shorter corresponding swing phase on the sound side (Murray et al., 1980). Walking is also characterized by the lateral movement of several body parts, especially during single support (standing on one leg) on the prosthetic side (1980). Moreover, the head and the neck deviate to the prosthetic side in stance and they also undergo vertical and forward movement with every step forward (1980). In addition, the shoulder on the sound side undergoes excessive flexion and the elbows exhibit excessive extension during walking (1980). The whole scene is accompanied by extensive hip hiking and low proprioception exhibited by an excessive prosthetic heel rise for ground clearance and a corresponding early plantar flexion on the sound side (1980). As explained by Murray and al. these abnormalities in gait are not only inevitable with the single axis knee, but also necessary for the deficiency in stability and control provided by this type of prosthetic knees (1980). It can thus be concluded that the single axis constant friction knee cannot be the prosthetic of choice when gait quality is an issue and thus, it cannot normalize the amputee’s mobility.

In addition, the multiple body movements mentioned above do not occur without a high metabolic cost. Therefore, besides the excessive energy required for the continuous control of the knee, amputees also spend additional energy in the abnormal body movements associated with their gait. In fact, our results show that the metabolic demand associated with the single axis constant friction knee can exceed that of a sound limb by 80% which make it only suitable for strong muscled amputees and highly tiring for the weak and the elderly (Mc Gimpsey, & Bradford, n.d.).

4.2.1.3. Conclusion

The single axis constant friction knee, though affordable and low maintenance, is a highly limiting knee and its use is accompanied by numerous disadvantages ranging from limited mobility, to limited speed of ambulation and a highly asymmetric gait, in addition to increased metabolic demand, making it suitable only for young and strong amputees.

    - The Jaipur Knee

The Jaipur knee is the result of the collaboration of the Jaipur foot organization and Stanford University and it has been selected to be one of the 50 best and most important inventions of the year 2009 (D-Rev Re:Motion Designs, What Gives 365, 2010)). At a manufacturing cost of $20, this polycentric knee matches and even excels $10000 titanium polycentric knees (The Jaipur knee Project I-Getting the Need Right, 2012). It owes its success to the numerous advantages associated with its use and to the minimum accompanying disadvantages.

4.2.2.1. Advantages

As shown by our results, the Jaipur knee is a polycentric knee that provides for an efficient swing control without the need for mechanical friction (Greene, 1983). The smoothness observed in the swing phase is due to the rotational movement of the center of the knee, characteristic of polycentric knees (1983). This movement regulates the acceleration and deceleration of the shin-foot complex according to the angle between the axis of the knee and the shin and results in an efficient swing of the leg that doesn’t require additional amputee muscle control (1983).

Moreover, the Jaipur knee also exhibits stance control achieved by a gentle locking of the knee when extended, thus preventing buckling problems and ensuring stability and security against falling (The Jaipur Knee Design Story, Vimeo).

These efficient swing and stance controls are accompanied by a smooth transition from stance to swing phase (Michal, n.d.). In addition, the knee offers 165° of flexibility, thus enabling kneeling and squatting, biking, walking on rough ground and of course, walking (D-Rev Re:Motion). With these features the Jaipur knee meets the needs of amputees in terms of movement possibilities, safety and stability, and enables them, even bilateral ones, to lead an independent, moderately active and productive life.

Another feature that places the Jaipur knee ahead of other conventional prosthetics, including the single axis knee, is the associated gait quality. The rotational center of the knee that relieves the amputee from additional muscle efforts, the stability in stance and the smooth swing control and movement of the leg contribute to an enhanced gait quality with reduced hip hiking and pelvic thrust (Greene, 1983, Radcliffe, 1977). The gait quality is also enhanced by the fact that unlike other knees, the prosthetic and the sound sides can be made to be equal and there is no need for shortening the prosthetic leg length (Greene, 1983). This is due to the polycentric nature of the knee that shortens the length of the leg as the knee goes from extension to flexion, just like a biological knee (1983). This feature is also important for efficient sitting with equal leg lengths on the prosthetic and sound sides (1983).

The enhanced gait quality is accompanied by a reduced metabolic demand. This reduction is achieved by the polycentric knee’s swing control capacities which decrease the effort required to control the knee, the efficient stance control that also reduces muscle efforts to control buckling, the shortening ability of the knee in flexion and the reduced demand on hip and pelvic muscles (Greene, 1983). The result is a reasonable gait with reduced fatigue. Thus, the Jaipur knee provides for an enhanced gait quality, reduced metabolic demand and fatigue, enabling above knee amputees to lead a more productive life.

Besides providing above knee amputees with a reasonable range of activity and mobility accompanied with stability, enhanced gait quality and reduced fatigue, the Jaipur knee, as shown by our results, exhibits a high degree of practicality of usage. Being made of a self lubricating polymer and high density polymer, it lubricates itself with every step and hence, requires no maintenance (USAID Science and Technology, 2010). Moreover, it is a durable knee and it has been bench-marked o have 3-5 years of life (D-Rev Re:Motion). Moreover, as with the Jaipur foot, fitting an above knee with the Jaipur knee including the whole prosthetic device, can be achieved in one single day, making it a practical and easy fitting (The Jaipur knee Project I-Getting the Need Right, 2012). Thus, the Jaipur knee is a highly practical and durable knee that requires no maintenance and can be fitted in record time, hence, another feature that puts it ahead of other above knee prosthetics.

Finally, when it comes to the advantages associated with the Jaipur knee, the $20 manufacturing cost stands out as the most prominent, especially that it is accompanied with the numerous advantages discussed above. This low cost makes it ideal for developing countries and NGOs, especially when considering the fact that 80% of the world amputees reside in developing countries and can’t afford prosthetic limbs (USAID Science and Technology, 2010).

4.2.2.2. Disadvantages

Our results indicate that the advantages accompanying the Jaipur knee outnumber by far the disadvantages. However, like all prosthetic limbs, especially conventional ones, disadvantages are inevitable. One such shortcoming is the linear response of the Jaipur knee to speed which makes it unable to accommodate variable walking speeds (Greene, 1983). This limitation is due to the polycentric nature of the knee (1983).

Another limitation is that being a mechanical knee, no matter how efficient in control and metabolic demand, it still depends on the amputee’s muscles for movement and thus depends on the amputee’s muscle conditions and his health (Michal, n.d.).

In addition, as with all mechanical prosthetics, there is a short delay between the sound leg and the prosthetic leg in the gait cycle, and the stance period on the prosthetic side is a bit longer than usual (Farber, & Jacobson, 1995). This short delay results in gait asymmetry and places additional strain on the sound leg. However, at low speeds, this effect is not highly significant and the gait quality is reasonably good.

However, as our reader might notice, these disadvantages are not major limiting factors, especially when compared to the major contributions of the Jaipur knee in enhancing mobility for above knee amputees, bearing in mind that above knee amputees have a very difficult time regaining mobility with conventional prosthetics.

4.2.2.3. Conclusion

The Jaipur knee, with its multiple advantages that include efficient swing and stance control, enhanced gait quality, reduced metabolic demand, and practicality of usage, can restore mobility for above knee amputees and give them a reasonably active and productive life, thus enhancing their quality of life.  In addition, when compared to the higher cost and highly limiting single axis constant friction knee, the Jaipur knee is certainly superior in all aspects, making it the prosthesis of choice for low income above knee amputees.

    - The BiOM AK: The X2 Bionic Knee

The BiOM AK is a robotic leg that combines the highly advanced X2 knee and the PowerFoot BiOM ankle (U.S. Department of Veteran affairs, 2012). As revealed by our results, the X2 knee, with its multiple sensors, accelerometer, gyroscope, microprocessors and hydraulic actuator is an intelligent and auto-adaptive knee that detects the position of the knee in space and the ground conditions and responds reflexively, thus providing for an effortless and thought free ambulation for above knee amputees (Edwards, 2011, Tan, 2012). The effortless mobility is completed with the power and intelligent system of the BiOM ankle (iWalk, 2012), resulting in a leg that facilitates and normalizes to a high degree the ambulation of above knee amputees, otherwise difficult.

4.2.3.1. Advantages

The use of the X2 is accompanied with numerous advantages. The intelligent swing control of the X2 knee and the auto- adaptive and weight independent stance control that enables the knee to lock in flexed positions provide for a wide variety of activities for above knee amputees, such as step-over step stair ascent and descent, successful obstacle cross over and even backward walking which are impossible with other knees (Edwards, 2011, Martin, Pollock, Hettinger, 2010). Moreover, the flexed locking feature of the knee provides for a secure mobility, because it prevents buckling not only in knee extension, but also in all knee positions (Edwards, 2011). In addition, the easily accessible 5 modes of functioning of the X2 knee add to the provided versatility of mobility, giving amputees the opportunity to pursue favored activities (2011). This high range of activities is accompanied by a high degree of stability provided by the reflexive adjustment of knee stiffness and position, making the X2 knee superior not only to the conventional knees discussed so far, but also to all existing prosthetic knees. Thus the X2 provides amputees with a high degree of mobility and enables them to lead an active life in all aspects.

Besides the numerously possible activities, the X2 knee also provides for instantaneous variation of walking speeds. As shown by our results, the knee, due to its auto-adaptive intelligent system, reacts to ground forces and adjusts the knee angle and knee stiffness reflexively and in real time, making the shift from slow walking to running not only possible, but also easy and swift, just like a biological leg and unlike the mechanical knees discussed previously (Orange Coast Prosthetics Inc., 2010-2012, Edwards, 2011).

Moreover, the variable speeds provided by the X2 knee, unlike conventional knees, are accompanied by a highly symmetrical gait even for bilateral amputees (OPC1media, 2012, Eyewitness News 5, 2012). This symmetry is further enhanced by the individual tuning and alignment facilities which are particularly useful for unilateral amputees (Edwards, 2011). While conventional knees exhibit increased gait asymmetry with increased mobility speeds, the X2 provides for a consistent and reasonably symmetrical gait which in turn reduces pain and discomfort at the level of the sound knee, the back and the hips, thus, enhancing amputees’ quality of life (Ostrovsky, 2009).

In addition to an enhanced versatility of mobility, accommodation of variable speeds, stability and enhanced gait quality, the X2 knee also provides for a reduced metabolic demand (U.S. Department of Veteran affairs, 2012). Although it is a passive device that doesn’t provide any power, its advanced control system and the hydraulic operation mechanism reduce the amputee’s need for additional muscle efforts usually used to control the knee as with conventional prosthetic knees (Progressive Orthotics and Prosthetics, 2012). Thus, through reducing metabolic demand, the X2 knee reduces the fatigue commonly experienced by above knee amputees.

The features discussed so far describe the capacity of the X2 knee alone. However, when combined with the BiOM ankle, the resulting prosthetic leg, the BiOM AK is a powerful prosthesis that enhances the performance of above knee amputees on all levels. The BiOM AK with the intelligent system of the ankle and the power provided by the foot with every step combined to the highly efficient X2 knee, not only increases the range of activities for amputee, but also enhances gait quality further and reduces metabolic demand (OPC1media, 2012). For example, the absence of power in the X2 knee is compensated for by the positive work provided by the powered ankle which propels the amputee forward with each step, thus greatly reducing the demand for amputee muscle work and reducing metabolic demand further (iWalk, 2012). In fact, the combination of the advanced controls of the X2 knee and the BiOM ankle and the provided power not only make for energy efficient and effortless ambulation, but also raise the gait quality to an amazing level (OPC1media, 2012). Moreover, the energy efficiency is also accompanied with a reduced cognitive demand and reduced psychological strain, in addition to reduced muscle strain and unnecessary fatigue (Hafner, 2007, U.S. Department of Veteran affairs, 2012). The amputee can thus walk without additional fatigue, without having to look down with every step and without worrying about stumbling. The high control system and energy efficiency of the BiOM AK come in contrast not only to the high energy demands of the single axis mechanical knee, but also to the efficient Jaipur knee and even to more advanced knees such as the C-Leg (Kaufman et al., 2008), making the BiOM AK superior in enhancing the quality of life of above knee amputees.

In terms of practicality of usage, the BiOM ankle proves to be practical in various aspects. The X2 knee and the BiOM ankle are easy to use because they operate reflexively and don’t need the amputee’s intervention and his conscious control for functioning (Edwards, 2011). Moreover, the 5 programmable modes of the X2 are easy to operate through a remote control device or by tapping the foot on the floor 3 times (2011). The 4-5 days rechargeable battery life of the X2 and the three long life rechargeable batteries of the ankle together with their LED lights systems to alert the user of battery charge status, further make the BiOM AK practical to use (Edwards, 2011, Herr, & Grabowski, 2011, Fierce Medical Devices, 2012).  Thus, although at first glance the BiOM AK may seem to be a complicated electronic device, especially if compared to simple conventional prosthetics like the single axis or Jaipur knee, its use as shown by our results is simple and easy and its practicality of usage further facilitates the difficult life of the above knee amputee.

4.2.3.2. Disadvantages

The use of the BiOM AK, like any other prosthetic limb, is also associated with disadvantages. However, these disadvantages as shown by our results are minor and do not cause significant limitations for amputees. One disadvantage associated with the BiOM AK is due to the fact that although the X2 can tolerate splashes of water, the ankle cannot and contact with water is damaging for the BiOM ankle, thus limiting the use of the BiOM AK to dry environments (Edwards, 2011, iWalk, 2012).

Another factor that can be considered limiting is the cost of the BiOM AK. While the X2 costs $32000, the BiOM AK costs around $80000 which can make it out of reach for low income amputees (Tan, 2012, Wynn, 2012). This cost can further increase due to required regular maintenance. If compared to the previously discussed conventional prosthetics, $80000 is viewed as a high cost for a prosthetic leg. However, if viewed in terms of the highly enhanced mobility, productivity and quality of life it provides, the cost factor becomes insignificant, bearing in mind that an economically productive amputee can even cover the cost of the prosthetic in the long run.

4.2.3.3. Conclusion

Above knee amputees face major difficulties in regaining mobility using conventional prosthetics. However, the BiOM AK, which combines the controls of the BiOM ankle and the X2 knee together with the power provided by the BiOM ankle, enables these amputees to achieve high levels and ranges of activity accompanied with security, reduced fatigue and reduced muscle and joint strains, in addition to a nearly normalized gait quality. Thus the BiOM AK not only enables above knee amputees to regain efficient mobility, but also enables them to regain control over their lives and become independent and active community members.

  - Upper Extremity Prosthetic Devices

Our evaluation of upper extremity prosthetic devices includes the conventional body- powered prosthetic arm, the bionic Luke arm coupled with Targeted Muscle Reinnervation and the recent Artificial Muscle Operated arm. The evaluation was performed against the five chosen criteria previously mentioned and the results obtained have led us to differentiate the advantages and disadvantages associated with each type. The following is a detailed discussion of our results.

    - Conventional or Body-Powered Arms

Body- powered prosthetic arms are conventional prosthetic arms that are made of a socket, a harness, a terminal device which can be a hook or a hand, and a cable system connecting the harness to the prosthetic joints (Kelly, 1994-2012). The design varies slightly with the degree of amputation, as does the scope of functionality of the arm (Wilson, 1963). The use of body powered prosthetic arms is associated with a series of advantages and disadvantages that vary with degrees of amputation.

4.3.1.1. Advantages

Body powered arms, as shown by our results, can prove to be useful in many ways, especially for below elbow amputees who have full control of the shoulder and elbow joints. The preferred terminal device being the hook because it offers visibility and thus, sensory feedback, is operated by a small tension applied to the cable connecting the hook to the harness by the amputee’s body muscles (Wilson, 1963, McKenzie, 1965). Below elbow amputees can thus engage in numerous activities including basic and daily life needs and lead an independent life (McKenzie, 1965). Moreover, task specific terminal devices enable them to lead a more active life, such as driving, hunting, skiing, eating, shaving, fishing, golfing and even playing musical instruments like the violin (McKenzie, 1965, Kelly, 1994-2012). The hand, though bulkier and more difficult to use than the hook, can be used for public appearance (Wilson, 1963). Thus, body powered prosthetic arms can prove advantageous to below elbow amputees and enable them to lead an active life.

For above elbow amputees, as shown by the results of our review of the literature, body powered arms can also provide for a reasonable range of activities with the aid of an elbow lock that locks the elbow after moving it to a desired position and flexing or extending it and an optional passive wrist unit that can rotate and be locked for increased functionality (Wilson, 1963). The operation of the elbow and the terminal device can be achieved using 1 or 2 cables, depending on the design (1963). The sensory feedback due to the increased visibility provided by the hook helps these amputees achieve a reasonable degree of independence in daily life in terms of basic needs operations.

The power used to move the body powered arm is provided by relative movements of sound body muscles such as the chest or shoulder muscles and the associated effort is considered reasonable for below elbow amputees (Wilson, 1963).

Below elbow amputees can also benefit from the moderate weight of the arm and from a loose socket that is secured by additional belts, making the prosthetic more comfortable (Wilson, 1963, Kelly, 1994-2012).

Other advantages associated with body powered arms include the durability of the device, the low maintenance requirements and the cost (Kelly, 1994-2012); at a maximum cost of $10000 body powered arms are the lowest functional prosthetic arms available (Mc Gimbsy, & Bradford, n.d.). These features make the body powered arms the most popular prosthetic arms and within the reach of most amputees.

4.3.1.2. Disadvantages

Unfortunately, our results show that the disadvantages and limitations in scope of functionality associated with body powered arms are numerous, especially for shoulder disarticulation cases. A major limitation is imposed by the control system of the arm which tends to increase in complication with the increase in the level of amputation. Consequently, above elbow amputees need to control the elbow and the terminal device individually, while shoulder amputees need to control the shoulder, the elbow and the terminal device, one joint at a time (McKenzie, 1965, Kuiken et al., 2009). These controls can involve 2 or 3 cables which prove to be demanding on the cognitive, muscle effort and energy demand levels and gets even more complicated for bilateral shoulder disarticulation cases (McKenzie, 1965). Moreover, the control of the arm in these cases needs long months of cognitively demanding training and the associated inefficiency in the scope of functionality compared to the excessive effort on the part of the amputee can be frustrating to a degree that some unilateral above elbow and shoulder amputees prefer switching to non-functional cosmetic arms (1965).

Another limiting factor associated with the use of body powered arms is the discomfort accompanying the use of the restricting and complicated harness. Harnesses differ in design depending on the degree of amputation and they are particularly complicated for shoulder amputees, especially women, because of the interference at breast level (Kelly, 1994-2012). Moreover, the weight of the arm which increases with increased amputation levels can be considered another discomfort factor (1994-2012).

In addition, and despite the tight and secure harness, additional straps and tight sockets, most amputees complain about the inefficiency of body powered arms because of sweating that leads to the slipping of the arm out of the socket easily, thus making it insecure (Adee, 2008).

Finally, body powered arms, with cables harnesses and hooks are very far from having a cosmetic appearance and this factor can be considered a significant disadvantage for public appearance (Harrison, 2008).

4.3.1.3. Conclusion

It can thus be concluded that body powered arms, though functional at a cost of $10000, are best tolerated and achieve maximum functionality in below elbow cases, giving below elbow amputees enough functionality to lead independent and active lives. However, these devices have major limitations in control, scope of functionality and comfort for above elbow and especially, shoulder amputees who need more efficient, practical and comfortable prosthetic arms to restore functionality and enhance their lives.

    - The Luke Arm

The Luke arm, as revealed by our review of the literature and media reviews, users’ comments and videos of interviews with Dean Kamen and DEKA engineers, is a highly advanced bionic arm that can be controlled mechanically, as well as by brain signals when coupled with targeted Muscle Reinnervation (TMR). It is a robotic arm that restores to a high degree the functionality of upper extremity amputees. The following is a detailed discussion of the advantages and disadvantages associated with the Luke arm.

4.3.2.1. Advantages

The Luke arm, with its many microprocessors, advanced sensory system, advanced socket, high degree of articulation, and its modular design that enables its customization according to amputation level, offers a multitude of advantages that can change the lives of amputees for the better.

With its 18° of freedom and the individual motors that power each joint separately, the Luke arm is close to the biological arm in articulation (Adee, 2008). As mentioned in our results, the shoulder, elbow, wrist and fingers are capable of movement in many directions and many planes enabling amputees a high degree of functionality (2008). From enabling shoulder amputees to reach behind their heads and perform lateral raises to inserting keys in a keyhole or picking an object as small as a grape and eating it, the Luke arm has proven to be capable of achieving an unprecedented level of operation (2008). It thus has the capacity to enable amputees regain control over their lives and live normal lives.

The sensory feedback provided by the sensor on the thumb connected to a vibrating motor that indicates the amount of grip pressure needed further enhances the efficiency of functionality of the Luke arm and enables amputees to handle delicate objects without damaging them (Adee, 2008). It thus does what a biological hand can do, and provides amputees with a normalized scope of functionality.

The control systems of the Luke arm constitute another advantageous feature that facilitates amputees’ lives. The arm as shown by our results is controlled by foot pads inserted in the amputee’s shoes and shoulder pads, both connected to the arm with cables (Adee, 2008). With simple and easy movements such as stepping on certain spots or moving the shoulder in certain ways, learned in 5-10 hours of training, amputees are able to control the am for maximum functionality such as inserting a key in a keyhole (2008). Moreover, when coupled with TMR, the control of the arm becomes even simpler. As explained in the results section, TMR involves the surgical rerouting of the nerves in the residual arm to a nearby new muscle where the nerves grow and branch, while this new location muscle is recognized by the brain as the missing arm (Kuiken et al., 2009). Brain signals involving arm movement orders reach the new location and lead to muscle contractions (2009). Electrodes on the corresponding skin surface pick up these signals and transmit them to the arm for operation (2009). When coupled with TMR, the muscle signals are amplified and result in an effortless and intuitively controlled arm movement (2009). Moreover, TMR enables amputees to operate different parts of the arm simultaneously and result in a fast and smooth movement of the arm, as is the case with biological arms (2009). The Luke arm, when coupled with TMR, loses its prosthetic attribute and becomes a real extension of the body. Thus, it not only replaces, but also restores the amputee’s arm. The Luke arm when coupled with TMR is especially useful for shoulder disarticulation cases, who due to inexistent residual muscles are unable to operate ordinary myoelectric arms that require strong and clear signals (Kuiken et al., 2005).

Another major advantage of the Luke arm is the associated comfort factor. With a mere 3.6Kg, which as shown by our results is equivalent to the weight of an average size woman’s arm, the Luke arm is comfortable to wear (Adee, 2008). Amputee comfort is further enhanced by the revolutionary socket system involving tiny balloons that inflate when the arm is in use to ensure a tight fit and deflate when relaxed, thus sparing amputees the discomfort of tight and sweaty sockets and slippery arms (CBS Interactive Staff, 2009). In addition, the arm is worn by strapping it across the chest and requires no restricting harnesses (Adee, 2008). Thus, the Luke arm with its comfortable weight, variable pressure socket and no harness features offers maximum comfort and practicality to amputees, and in doing so it comes closer to the biological arm. Moreover, the comfort factor is further enhanced by the fact that the arm requires no muscle work and all the work required to move it is supplied by the motors (2008). These features are highly appreciated by above elbow, unilateral and bilateral shoulder amputees who as mentioned usually give up on conventional body powered prosthetic arms and also face major difficulties with the use of ordinary myoelectric arms (Dailami, 2002, Kuiken, 2005). The Luke arm thus, with its multiple advantages, enables all amputees to regain arm functionality and lead a normal and active life.

4.3.2.2. Disadvantages

The disadvantages associated with the use of the Luke arm are minor and do not involve its functionality. One disadvantage is associated with TMR and the period of 0.5-1.5 years required for the reintegration of the redirected nerves in the new location, in addition to a long training period to achieve an efficient brain-arm link (Ottobock , 2012b, Kuiken, 2005). Another disadvantage is the foot pad control mechanism, which as stated by one amputee makes the operation of the arm possible only while standing and not while walking (Gould, 2012). In addition, the mechanical noise associated with the movement of the arm can be an inconvenience for self conscious amputees.

The cost of the Luke arm, on the other hand can prove to be a significant limiting factor for low income amputees, especially when not covered partially or totally by insurance. At a cost slightly higher than$100000, which is the usual cost of most myoelectric arms, the Luke arm can be considered high cost (Saenz, 2009). Moreover, when coupled with TMR this cost rises by $300000 and this can make it beyond the reach of most amputees, thus limiting its use (Ryerson University, 2011). However, as clarified in our discussion of advanced lower limb prosthetics, this high cost becomes less significant when viewed from the normalized productivity perspective.

4.3.2.3. Conclusion

The multiple advantages of the Luke arm make it beyond compare and place it ahead of all prosthetic arms, especially when coupled with TMR. The Luke arm not only moves beyond the limited functionality, great discomfort and frustrating controls of conventional arms, but also surpasses the ordinary myoelectric arms in its efficiency and thus stands out as the prosthetic arm of choice for all amputees, giving them a new beginning and restoring their quality of life.

    - The Artificial Muscle Operated Arm (AMO)

The Artificial Muscle Operated arm is an advanced and brain controlled prosthetic arm that has won its designers many awards (CBC, 2011). It owes its success to its reasonable cost and its simple yet efficient design that comprises a headset containing sensors and an arm containing a microprocessor and artificial muscles that are powered by a pneumatic pump connected to a compressible air tank (Ryerson University, 2011). This simple and intuitively controlled arm, as shown by our research results comes with significant advantages as well as a few disadvantages of varying degrees of limitation. The following is a detailed discussion of our results.

4.3.3.1. Advantages

The most significant advantage of the AMO arm is that it is a brain controlled prosthetic arm that requires no invasive surgery (Ryerson University, 2011). The sensors worn superficially on the skull have the capacity of recording and transmitting brain signals efficiently (Hein, 2011). These signals are then interpreted by the microprocessor in the arm which sends out a matching signal to start the desired operation of the arm (Ryerson University, 2011). Thus, the simple operation mode of the AMO arm is an advantageous feature, because it is independent of body and muscle contraction capabilities and this feature makes the use of the arm easy and possible for all levels of amputations and without the cost of complicated surgery, especially given the fact that TMR is still unavailable in most countries.

Another equally important feature of the AMO arm is that it is the lowest cost, brain controlled prosthetic arm to date (Hayes, 2011). While myoelectric arms cost $100000 and their efficiency depends on the amputee’s capacity of producing high EMGs, and TMR costs $300000 and is only available in very few hospitals and very few countries around the world, the AMO arm at only $25000, stands out as a feasible solution for low income amputees, especially with the promised future design modifications that will further enhance its articulation capacities and its practicality features (Ryerson University, 2011).

In terms of functionality, the AMO arm provides for the basic functions required to lead a reasonably independent life. As shown by our results, the arm is capable of various articulations. It can raise and lower the arm at shoulder level, flex the elbow, rotate the arm and close the fingers to hold an object (CBC News, 2011). The scope of functionality of the AMO arm in its current state can’t match that of the Luke arm, but it is certainly superior to that of body powered arms that are especially inconvenient for high levels of amputation.

Moreover, the electrodes in the headset have the capacity to record and transmit brain signals with enough efficiency, and the signals picked up from other sources are negligible (Sabah N., Personal Communication, December 18, 2012). This efficiency further adds to the advantages of the AMO arm and facilitates the amputee’s operation of the arm.

In terms of practicality of usage, the AMO not only provides for basic arm operations, but is also a powerful arm and the artificial muscles in the arm can lift up to 200lb of loads (Discovery Channel, 2011). Moreover, as shown by our research, it does not require an external energy source. The power source of the AMO is the compressed air in the associated air tank and this source can last 16 hours of usage, thus adding to its practicality of usage (2011).

The AMO arm is also advantageous in terms of comfort. The comfort associated with its use is the result of its light weight due to the absence of motors and complicated circuits. The absence of annoying and restricting harnesses is another major comfort factor (Hayes, 2011). The arm is simply worn strapped across the chest, making its use easy and comfortable, thus enhancing the lives of amputees without any undue discomfort (2011). In addition, our results indicate that the headset is noiseless, does not cause discomfort and most importantly, does not harm the brain in any way (Sabah N., Personal Communication, December 18, 2012).

4.3.3.2. Disadvantages

The AMO in its current design specifications does not come without disadvantages. One significant disadvantage is the limited degree of articulation, especially at the level of the individual articulation of the fingers, in addition to the lack of sensory feedback (Ryerson, 2011). Thus, the AMO at its current stage cannot provide for delicate and precise operations such as picking up a raisin or inserting a key in a keyhole and unlocking a door (2011). However, as promised by its designers, future improvement of the design will solve these problems (2011).

Moreover, the AMO, like all brain controlled arms, can’t be used without sufficient training. Although the basic operations can be achieved in less than 10 minutes, more complicated operations require a few weeks of serious training (Discovery Channel, 2011). Our results suggest that further training is also needed to overcome the limitations or problems associated with the wireless transmission, since superficially placed electrodes can’t be 100% efficient in recording all the emitted brain signals and training is required to produce the required intensity of signals, so as to overcome to a reasonable degree this limitation (Sabah N., Personal Communication, December 18, 2012).

Other minor disadvantages include the use of the headset which can be annoying or not cosmetic for some amputees, especially in public. Another such minor disadvantage is the compressible air tank which, in the current design, has to be carried in a pocket or on the body (Ryerson University, 2011). However, future designs will include the tank inside the arm, thus making the arm more practical to use (2011). Our results suggest one last disadvantage associated with the wireless connection which is the fact that the AMO arm can be subject to possible electromagnetic interference which can disrupt the wireless transmission occasionally and lead to operation problems (Sabah N., Personal Communication, December 18, 2012). However, as specified, these limitations are minor and do not cause any long term or serious operation inefficiency.

4.3.3.3. Conclusion

The AMO arm, despite the limitations in degrees of freedom and the lack of sensory feedback associated with its current design, in addition to the minor disadvantages previously discussed, is a functional, efficient and comfortable arm that can facilitate the life of an arm amputee and aid him to lead an independent and comfortable life using only his brain signals. Though not as efficient as the highly advanced Luke arm, at $25000 the AMO arm relieves amputees from the limited functionality of body powered arms and spares them the discomfort associated with harnesses, controlling cables, tight and sweaty sockets and increased muscle work. It can thus be concluded that the AMO is a feasible and good choice in case the Luke arm is unattainable for financial reasons.

- Conclusion and Recommendations

From what has been presented so far, we can conclude that the benefits of the advanced line of prosthetics analyzed in this report which includes the BiOM ankle for below knee cases, the BiOM AK for above knee cases and the Luke arm coupled with TMR for all upper extremity cases, outweigh their high costs. Moreover, if the minimum loss of $107060 for each economically unproductive amputee over a 20 year period is taken into consideration, the cost factor becomes even less significant in terms of limiting the feasibility of such high performance prosthetic devices. Productive amputees are more capable of covering their prosthetics expenses than amputees using limiting prosthetics, and the advanced prosthetics discussed in this report are capable of restoring their productivity, hence ensure their rise on the economic level.

Moreover, our analysis of the AMO arm has led us to conclude that advanced and brain controlled prosthetics can be made simpler, cheaper, and most importantly, locally made so as to help a maximum number of amputees at minimum cost. Consequently, in order to take our case further, we recommend the following actions:

- A more thorough cost-benefit analysis of the BiOM ankle, the BiOM AK and the Luke arm, in order to determine long term benefits of these devices on all levels, with the aim of increasing funding resources for fitting a maximum number of amputees with these advanced prosthetics.

- A thorough study of the technology involved in advanced prosthetics, with the aim of producing locally made advanced prosthetic limbs that normalize limb operation.

- A thorough study of the possible funding for such advanced prosthetics, in terms of project and application.

Finally, we can conclude that advanced prosthetic limbs are a necessity for today’s amputee and no effort should be spared in order to aid Lebanese victims to have access to these devices. We believe that with initiatives like that of BLOM Bank which offers 50% of its credit card profits to help amputees, the fund raising efforts of NGOs and beneficiary organizations and perhaps, with long term and low or no interest loans for amputees, these high cost prosthetics can be within the reach of a great number of Lebanese landmine and UXO surviving victims.

References:

Adalarasu K., (2011), Comparison on Jaipur, SACH and Madras Foot A Psychophysiological Study, International Journal of Advanced Engineering Sciences and Technologies, Volume 4, Issue 1, 187-192. Available: http://ijaest.iserp.org/archieves/10-My1-15-11/Vol-No.6-Issue-No.2/5.IJAEST-Vol-No-6-Issue-No-2-Comparison-on-Jaipur,-SACH-and-Madras-Foot-187-192.pdf

Adee, S., (2008, February), Dean Kamen's "Luke Arm" Prosthesis Readies for Clinical Trials, IEEE Spectrum. Available: http://spectrum.ieee.org/biomedical/bionics/dean-kamens-luke-arm-prosthesis-readies-for-clinical-trials

# Adee S. (2008), Dean Kamen's Robotic "Luke" Arm, IEEE Spectrum Online. Retrieved from:  http://www.youtube.com/watch?v=R0_mLumx-6Y

Au S.K., Herr H., (2008, September), Powered Ankle-Foot Prosthesis the Importance of Series and Parallel Motor Elasticity, IEEE Robotics and Automation Magazine, Volume 15, Issue 3, 52-59. Available: http://ieeexplore.ieee.org.ezproxy.aub.edu.lb/search/searchresult.jsp?newsearch=true&queryText=the+importance+of+series+and+parallel+motor+elasticity&x=14&y=20

Au S. K., Herr H., Weber J., Martinez-Villalpando E. C., (2007), Powered Ankle-Foot Prosthesis for the Improvement of Amputee Ambulation, Engineering in Medicine and Biology Society, 2007, EMBS 2007, 29th Annual International Conference of the IEEE, pages 3020-3026. Available: http://ieeexplore.ieee.org.ezproxy.aub.edu.lb/stamp/stamp.jsp?tp=&arnumber=4352965

Background on Landmine Problem in Lebanon, Available: http://www.balamand.edu.lb/English/FHSCAProgramFrame.asp?id=9624&fid=653&PageName=Background%20on%20Landmine%20Problem%20in%20Lebanon

Barr A., Lohmann Siegel K., Danoff J., McGarvey III Ch., Tomasko A, Sable I, and Stanhope J., (1992, May), Biomechanical Comparison of the Energy- Storing Capabilities of SACH and Carbon Copy II Prosthetic Feet During the Stance Phase of Gait in a Person with Below-knee Amputation, Journal of the American Physical Therapy Association, Volume 72, No.5, 344-354. Available: http://www.physicaltherapyjournal.com/content/72/5/344.full.pdf

# Blough D.K.,  Hubbard Sh.,  McFarland L.V., Smith D.G.,  Gambel J.M.,   Reiber G.E., ( 2010, June 15), Prosthetic cost projections for service members with major limb loss from Vietnam and OIF/OEF, Journal of Rehabilitation Research and Development, Volume 47 Number 4, 2010,   387 — 402. Available: http://www.rehab.research.va.gov/jour/10/474/blough.html

Burger N., Analysis of Amputee Gait, Atlas of Limb Prosthetics: Surgical, Prosthetic and Rehabilitation Principles, Chapter 14. Available: Digital Resource Foundation for Orthotics and Prosthetics Community, http://www.oandplibrary.org/alp/chap14-01.asp

Carpenter M., Hunter C., Rheaume, D., (2008, April 24), Testing and Analysis of Low Cost Prosthetic Feet, Worcester Polytechnic Institute. Available: http://ebookbrowse.com/testing-and-analysis-of-low-cost-prosthetic-feet-pdf-d24619968

CBC News, (2011, May 17), Thought Controlled Prosthetic Vies for Prize, CBC News Technology and Science. Available: http://www.cbc.ca/news/technology/story/2011/05/17/technology-prosthetic-arm-prwata-caires.html

Clements, I.P., (1998-2012), How Prosthetic Limbs Work, How Stuff Works, p.5. Available, http://science.howstuffworks.com/prosthetic-limb4.htm

Dailami J.,,(2002,May 4), The Myoelectric Arm It’s Electrifying, Illumin University of Southern California School of Engineering, Volume XIII, Issue III,2012, December 7.Available: http://illumin.usc.edu/111/the-myoelectric-arm-it39s-electrifying/1/

## Davies and Starr, (2009), The 50 Best Inventions of 2009, The $20 Knee, Time Magazine. Available:

http://www.time.com/time/specials/packages/article/0,28804,1934027_1934003_1933963,00.html

# Discovery Channel, (2011, May 6), Bionik Labs on Discovery Channel - Michal Prywata: Retrieved from: http://www.youtube.com/watch?v=vQvsMJ9Hw4Y

D-Rev Re:Motion Designs, Jaipur Knee: A prosthetic knee joint for extreme affordability. Available: http://www.d-rev.org/assets/JaipurKnee_Overview.pdf

Dr Pramod Sethi, Inventor of Jaipur Foot, Dead, The Indian Express, (2008, January 6), Available: http://www.indianexpress.com/news/dr-pramod-sethi-inventor-of-jaipur-foot-dead/258475

# Dupes B., ( 2008, September 18), Prosthetic Knee Systems, Military in-Step. Available: http://www.amputee-coalition.org/military-instep/knees.html

Edelstein J., (1988, December), Prosthetic Feet State of the Art, Journal of the American Physical Therapy Association, Volume 68, No. 12, 1874-1881. Available: http://ptjournal.apta.org/content/68/12/1874.full.pdf

Edwards M., (2011, July 28), Webinars on Products and Services from Ottobock: Microprocessor Knees Family of Knees. Available: http://academy.ottobockus.com/training/product-webinars.aspx

Eilenberg M. F., Geyer H., Herr H., (2010), Control of a Powered Ankle-Foot Prosthesis Based on a Neuromuscular Model, Neural Systems and Rehabilitation Engineering, IEEE Transactions, Volume 18, Issue 2, 164-173, Available: http://ieeexplore.ieee.org.ezproxy.aub.edu.lb/stamp/stamp.jsp?tp=&arnumber=5378631

E-Mine, (2009, September), Victim Assistance, Electronic Mine Information. Available: http://www.mineaction.org/overview.asp?o=18

Eyewitness News 5, (2012, June 26), Heath Calhoun Fit with X2 Knees. Retrieved from: http://www.youtube.com/watch?v=TlHAuhs9k84

## Fairley M., (2007, December), Targeted Muscle Reinnervation: The Future Is Now, O&P Edge. Available: http://www.oandp.com/articles/2007-12_01.asp

Farber B., Jacobson J. (1995, November)An Above-Knee Prosthesis with a System of Energy Recovery:A technical note, Journal of Rehabilitation Research and Development, Volume 32 Number 4, pages 337-348. Available: http://www.rehab.research.va.gov/jour/95/32/4/pdf/farber2.pdf

Gailey R., (2005), Functional Value of Prosthetic Foot/Ankle Systems to the Amputee, American Academy of Orthotists and Prosthetists, JPO, Volume 17, No. 4S, 39-41. Available: http://www.oandp.org/jpo/library/2005_04S_039.asp

Goodier R., (2011, March 6), Jaipur: A Low-Cost Prosthetic Gives Legs to Amputees in the Developing World, Available: http://www.google.com.lb/search?q=Jaipur%3A+a+low-cost+prosthetic+gives+legs+to+amputee+in+the+developing+world&oq=Jaipur%3A+a+low-cost+prosthetic+gives+legs+to+amputee+in+the+developing+world&sugexp=chrome,mod=0&sourceid=chrome&ie=UTF-8

Gordon E. and Ardizzone J., (1960, March 01), Clinical Experiences with S.A.C.H. Foot Prosthesis, The Journal of Bone and Joint Surgery, Volume 42, Issue 2, 226-234. Available: http://jbjs.org/article.aspx?articleid=13255

Greene M.,(1983), Four Bar Linkage Knee Analysis. Available: http://www.oandplibrary.org/op/pdf/1983_01_015.pdf

Hafner B.J., Willingham L.L., Buell N.C., Allyn K.J., Smith D. G.,(2004-2007), Evaluation of Function, Performance, and Preference as Transfemoral Amputees Transition From Mechanical to Microprocessor Control of the Prosthetic Knee, Archives of Physical Medicine and Rehabilitation, Volume 88, Issue 2, February 2007, Pages 207–217. Available: http://www.sciencedirect.com.ezproxy.aub.edu.lb/science/article/pii/S0003999306014808

Gould J., 2012, March 10, New prosthetic arm born from ‘Star Wars’, Marine Corps Times. Available: http://marinecorpstimes.com/news/2012/03/marine-luke-arm-prosthetic-born-from-star-wars-031012w/

Harrison K., (2008, April), A closer look at Componentry Upper limb prosthetics, Limbs 4 Life, Available: http://www.limbs4life.com/content/articles/componentry/a%20closer%20look%20at%20componentry_upper_limbs%20_web.pdf

# Hayes M., (2011, May 6), Bionik Labs on CHCH - Michal Prywata, CHCH Health News. Retrieved from: http://www.youtube.com/watch?v=uNmlvGURBZk

Hein T., (2011, August 15) We Have the Technology…, Canadian Manufacturing: Design Engineering. Available: http://www.canadianmanufacturing.com/design-engineering/case-studies/we-have-the-technology-39843

Hill, G., A Closer Look at Componentry Feet, Limbs4Life, Available: http://www.limbs4life.com/content/articles/componentry/A%20CLOSER%20LOOK%20AT%20COMPONENTRY-Feet.pdf

Herr H., Grabowski A., (2011), Bionic Ankle-Foot Prosthesis Normalizes Walking Gait for Persons with Leg Amputation, Proceedings of the Royal Society, Available: http://www.google.com/url?sa=t&rct=j&q=&esrc=s&source=web&cd=2&ved=0CDsQFjAB&url=http%3A%2F%2Frspb.royalsocietypublishing.org%2Fcontent%2Fearly%2F2011%2F07%2F07%2Frspb.2011.1194.full.pdf&ei=5TipUKzgJcyM4gSlu4CQCQ&usg=AFQjCNHJkc0DDAgHJTzgEZ3PFibLFecDKw

ICRC, (2005, June), Caring for Landmine Victims, Available: http://www.icrc.org/eng/assets/files/other/caring-landmine-victims-0863.pdf

# iWalk Enhances Bionics Experience with New Personal Power Package for BiOM® Ankle System, (2012, October 24), Fierce Medical Devices, Available: http://www.fiercemedicaldevices.com/press-releases/iwalk-enhances-bionics-experience-new-personal-power-package-biom-ankle-sys

iWalk, (2012), http://www.iwalkpro.com/InTheNews.html

iWalk, (2012), http://www.iwalkpro.com/InTheNews/faq.htm

Jaipur Foot, (2007), Bhagwan Mahaveer Viklang Sahayata Samiti. Available: http://www.jaipurfoot.org/

Kapp S, Cummings D, Transtibial Amputation: Prosthetic Management, Chapter 18B-Atlas Limb Prosthetics: Surgical, Prosthetic Rehabilitation Principles, Digital Resource Foundation for the Orthotics and Prosthetics Community. Available: http://www.oandplibrary.org/alp/chap18-02.asp

Kaufman K.R., Levine J.A., Brey R.H., McCrady S.K., Padgett D.J., Joyner M.J., (2008), Energy Expenditure and Activity of Transfemoral Amputees Using Mechanical and Microprocessor Controlled Prosthetic Knees, Archives of Physical Medicine and Rehabilitation

Volume 89, Issue 7, July 2008, Pages 1380–1385. Retrieved: http://www.sciencedirect.com.ezproxy.aub.edu.lb/science/article/pii/S0003999308002785

Keefe B., Bionic Soldiers, (2011, July 25), abc 9 News. Retrieved December 22, 2012 from: http://www.youtube.com/watch?v=52tFxg0O_bs

Kelly B. M., (1994-2012), Upper Limb Prosthetics, Medscape Reference by WebMD LLC. Available, http://emedicine.medscape.com/article/317234-overview#a17

# Knapp A., (2012, April 4), iWalk Has Created The First Truly Bionic Foot, Forbes. Available: http://www.forbes.com/sites/alexknapp/2012/04/19/iwalk-has-created-the-first-truly-bionic-foot/

## Kroeker K.L., (2011, April), Engineering Sensation in Artificial Limbs, Communications of the ACM, Volume 54, Number 4, 16-17.

## Kuiken T.A., Dumanian G.A., Lipschutz R.D., Miller L.A., Stubblefield K.A., (2005, March 16-19), Transferring Muscle Reinnervation for Improved Myoelectric Prosthesis Control, Neural Engineering 2005, Conference Proceedings 2nd Interaction IEEE EMBS Conference, Arlington, Virginia, pp. 369-399. Available: http://ieeexplore.ieee.org.ezproxy.aub.edu.lb/stamp/stamp.jsp?tp=&arnumber=1419642

# Kuiken T., Li G., Lock B.A., Lipschutz R.D., Miller L.A., Stubblefield K.A., Englehart K.B., (2009, February 11),Targeted Muscle Reinnervation for Real-time Myoelectric Control of Multifunction Artificial Arms, The journal of the American Medical Association, Volume 301, Number 6. Available: http://jama.jamanetwork.com/article.aspx?articleid=183371

# Layton J., How can someone control a machine with her thoughts? Howstuffworks, Available: http://science.howstuffworks.com/bionic-arm.htm

Lee D., (2012, September 10), Bionic Revolution: The tech getting disabled people into work, Available: www.bbc.co.uk/news/business-19547954

Lee M., Moroz A., (2009), Physical Therapy, The Merck Manual for Health Care Professionals. Available: http://www.merckmanuals.com/professional/special_subjects/rehabilitation/physical_therapy_pt.html

Mancinelli CH., Patriti B., Tropea P. Greenwald R., Casler R., Herr H., Bonato P., (2011, August 30-September 3), 33rd Annual International Conference of the IEEE EMBS Boston, Massachusetts. Available: http://ieeexplore.ieee.org.ezproxy.aub.edu.lb/stamp/stamp.jsp?tp=&arnumber=6092035

Marinakis G., (2004, July/August), Interlimb Symmetry of Traumatic Unilateral Transtibial Amputees Wearing Two Different Prosthetic Feet in the Early Rehabilitation Stage, Journal of Rehabilitation Research and Development, Volume 41, No. 4, 581-590. Available: http://www.rehab.research.va.gov/jour/04/41/4/pdf/Marinakis.pdf

## Martin J., Pollock A., Hettinger J.,(2010), Microprocessor Lower Limb Prosthetics: Review of Current State of the Art, American Academy of Orthotists and Prosthetists,  JPO, Volume 22, Number 3, pp. 183-193. Available: http://www.oandp.org/jpo/library/2010_03_183.asp

Martin P., (2011, January), Return to Combat, The O&P Edge. Available: http://www.oandp.com/articles/2011-01_01.asp

Mathur M. K., Jaipur Artificial Limbs. Available: http://www.jaipurfoot.org/pdf/Jaipur-Artificial-Limbs.pdf

Mc Gimpsey G. and Bradford T. C., Limb Prosthetics Services and Devices Critical Unmet Need: Market Analysis White Paper, Bioengineering Institute Center for Neuroprosthetics,

Worcester Polytechnic Institution. Available: http://www.nist.gov/tip/wp/pswp/upload/239_limb_prosthetics_services_devices.pdf

Mc Kenzie D. S., (1965, August), The Clinical Application of Externally Powered Artificial Arms, The Journal of Bone and Joint Surgery, Volume 47B, Number 3. Available: http://www.bjj.boneandjoint.org.uk/content/47-B/3/399.full.pdf

Michael J., Are you Ready for a Prosthesis, Prosthetic Knees, First Step,46-49. Available: http://www.abcamputee.com/readyforpros.pdf

Murray P. et al.,(Fall 1980)  Gait Patterns of Above-Knee Amputees Using Constant-Friction Knee Components, Bulletin of Prosthetics Research BPR, 10-34, Volume 17, Number 2,Pages 35-45. Available: http://www.rehab.research.va.gov/jour/80/17/2/murray.pdf

Nobbe Orthopedics Inc., (2001), Best Feet Forward, Prosthetic and Orthopedic Update, A Publication of Nobbe Orthopedics, Number 30. Available: http://nobbeorthopedics.com/ns_pdf/nobbe_news_letter_30.pdf

# OPC1media, (2012, March 6), X2 + BiOM Foot = Astounding Results, Orthotic Prosthetic Center Inc. Retrieved from: http://www.youtube.com/watch?v=BJx9pFa2Fec

Orange Coast Prosthetics Inc., (2010-2012), The New Genium/X2. Bionic Prosthetic System. Available:  http://ocprosthetics.com/The_New_Genium___X2.html

Orthotic and Prosthetic Appliances: Billing Codes and Reimbursement Rates-Prosthetics, (2012, January). Available: http://www.google.com.lb/search?q=orthotic+and+prosthetic+appliances%3Abilling+codes+and-medi-cal&oq=orthotic+and+prosthetic+appliances%3Abilling+codes+and-medi-cal&sugexp=chrome,mod=4&sourceid=chrome&ie=UTF-8#hl=en&sugexp=les%3B&gs_nf=3&pq=orthotic%20and%20prosthetic%20appliances%3Abilling%20codes%20and-medi-cal&cp=76&gs_id=1o&xhr=t&q=orthotic+and+prosthetic+appliances:billing+codes+and-medi-cal+september+2012&pf=p&safe=off&tbo=d&sclient=psy-ab&oq=orthotic+and+prosthetic+appliances:billing+codes+and-medi-cal+september+2012&gs_l=&pbx=1&bav=on.2,or.r_gc.r_pw.r_qf.&fp=b15c94e84fef4736&bpcl=38897761&biw=1241&bih=606

# Ostrovsky G.,(2009, December 14), X2 from Otto Bock Brings Smarts to Artificial Legs, MedGadget. Available: http://www.medgadget.com/2009/12/x2_from_otto_bock_brings_smarts_to_artificial_legs.html

Ottobock, (2012), SACH. Available: http://www.ottobock.com/cps/rde/xchg/ob_com_en/hs.xsl/5767.html

Otto Bock, (2012), Targeted Muscle Reinnervation, Ottobock Health Care, Available: http://www.ottobock.com/cps/rde/xbcr/ob_com_en/646D385-GB-03-1006w.pdf

Prahald C. K., (2003, December 12), Jaipur Foot: Challenging Convention, Michigan Business School. Available: http://www.bus.umich.edu/FacultyResearch/ResearchCenters/ProgramsPartnerships/IT-Champions/JaipurFoot.pdf

Progressive Orthotics and Prosthetics, (2012), Genium is Here. Available: http://www.progoandp.com/genium-is-available.html

Prokaza J., (2012, February), Has the Bionic Man Arrived? Computer Shopper, Issue 288, page 110. Available: http://www.google.com/url?sa=t&rct=j&q=&esrc=s&source=web&cd=1&ved=0CDAQFjAA&url=http%3A%2F%2Fwww.iwalk.com%2Fdocs%2FFEATURE%2520THE%2520BIONIC%2520MAN_Computer%2520Shopper%2520Magazine.pdf&ei=hESpUKqTN-mm4gSY74GwDw&usg=AFQjCNGWGBEKEtVEQkIfFNvZpLym5zGWAA

## Radcliffe C. W., (1977), The Knud Jansen Lecture: Above-knee prosthetics, O&P Library, POI, Volume 1, Number 3, 146-160. Available:

http://www.oandplibrary.org/poi/1977_03_146.asp

Rosenberg T., (2011, June 28), Keeping Artificial Limbs Lo-Cost, and High-Quality, The NewYork Times, Opinionator. Available: http://opinionator.blogs.nytimes.com/2011/06/28/keeping-artificial-limbs-low-cost-and-high-quality/

### Ryerson University, (2011, March 29), Ryerson students invent breakthrough brain-controlled prosthetic arm, Ryerson University News and Events. Available: http://www.ryerson.ca/news/media/General_Public/20110329_RN_CairesPr.html

SACH Foot, Sheck and Siress. Available: http://www.scheckandsiress.com/care_maintenance.as

Saenz A., (2009, January 12), DEKA’s Luke Arm in Clinical Trials, Is it the Future of Prosthetics?, Singularity Hub, Available: http://singularityhub.com/2009/12/01/dekas-luke-arm-in-clinical-trials-is-it-the-future-of-prosthetics-video/

Sebastian S., (2012, March 6), Jaipur Foot Gets an American Facelift, The Hindu. Available: http://www.thehindu.com/todays-paper/tp-national/article2965082.ece

Sethi P. K., Udawat M. P., Kasliwal S. C. (1978), Vulcanized Rubber Foot for Lowe Limb Amputees, Prosthetics and Orthotics International, No. 2, 125-136.  Available: www.oandplibrary.org/poi/pdf/1978_03_125.pdf

Strickland E., (2012, August 30, A Robot Ankle for Amputees, IEEE Spectrum. Available: http://spectrum.ieee.org/video/biomedical/bionics/a-robot-ankle-for-amputees

Tan M., (2012, March, 14), DoD Says Amputations Reached Wartime High, Army Times, Available: http://www.armytimes.com/news/2012/03/army-amputations-reach-war-time-high-031212w/

The Jaipur Knee Design Story, Basset Inspiration Series, Retrieved on November 27, 2012, from http://vimeo.com/26994102

The Jaipur Knee Project I-Getting the Need Right, (2012, August), Global Health Innovation Blog. Available: http://stanfordglobalhealth.com/2012/08/03/the-jaipurknee-project-i-getting-the-need-right/

# The JaipurKnee Project II – Scaling Up the Business,(2012, August),Global Health Innovation Blog Available: http://stanfordglobalhealth.com/2012/08/03/the-jaipurknee-project-ii-scaling-up-the-business/

# Trout Ch., (2011, April 5), AMO Arm pneumatic prosthetic does mind-control on the cheap, Engadget. Available:

http://www.engadget.com/2011/04/05/amo-arm-pneumatic-prosthetic-does-mind-control-on-the-cheap/

United States, Department of Veteran Affairs, (2012, July 18), VA Delivers Cutting-Edge Bionics to Above-Knee Amputees, Available: http://www.va.gov/health/NewsFeatures/20120719a.asp

USAID Science and Technology, (2010, October 1), Jaipur Knee. Available: http://transition.usaid.gov/scitech/jk.html

Versluys R., Desomer A., Lenaerts G., Beyl P., Van Damme M., Vanderborght B., Vanderniepen I., Van der Perre G., and Lefeber D., (2008, October 19-22), From Conventional Prosthetic Feet to Bionic Feet: A Review Study, Proceedings of the 2nd Biennial IEEE/RAS-EMBS International Conference on Biomedical Robotics and Biomechatronics, Scottsdale, AZ, USA. Available:

http://ieeexplore.ieee.org.ezproxy.aub.edu.lb/search/searchresult.jsp?newsearch=true&queryText=From+Conventional+Prosthetic+Feet+to+Bionic+Feet%3A+A+Review+Study&x=44&y=14

Wallach J., Sabiel E., (1970), Control Mechanism Performance Criteria for an Above-Knee Leg Prosthesis, Journal of Biomechanics, Volume 3, Issue 1, January 1970, Pages 87–88, IN7, 89–97. Available: http://www.sciencedirect.com.ezproxy.aub.edu.lb/science/article/pii/0021929070900539

What Gives 365, (2010, MARCH 25), 5 pieces of plastic + 4 nuts and bolts = The ability to walk again. Available: http://whatgives365.wordpress.com/2010/03/25/5-pieces-of-plastic-4-nuts-and-bolts-the-ability-to-walk-again/

Wilson B. A. Jr., (1963), Limb Prosthetics Today, O & P Library, Artificial Limbs, Volume 7, Number 2, 1-42. Available: http://www.oandplibrary.org/al/1963_02_001.asp

##### Wilson B.A., (1968), Recent Advances in Above-Knee Prosthetics, Artificial Limbs, O&P Library, 1968, Volume 12, Number 2, 1-27. Available:

http://www.oandplibrary.org/al/1968_02_001.asp

Wynn K., (2012, April 26), Local Vets Chosen for 1st Bionic Ankle, Dayton Daily News. Available: http://www.daytondailynews.com/news/news/local/local-vets-chosen-for-1st-bionic-ankle-system/nMz7s/

Yammine M. J., (2007, October 30), The Impact of Landmines and Unexploded Ordnance on the Human and Land Resources of South Lebanon. Available: https://scholarworks.aub.edu.lb/bitstream/handle/10938/7551/st-5011.pdf?sequence=1

Recommended Short Videos

A Robot Ankle for Amputees: http://www.youtube.com/watch?v=HhSVqsHzRl4

BiOM Ankle Prosthetic: http://www.youtube.com/watch?v=-MsRrr_y_Sg

Power Foot BiOM by iWalk: http://www.youtube.com/watch?v=0tsAmiz9w1g

# iWalk's BiOM bionic ankle for lower-limb amputees:

http://www.youtube.com/watch?v=WlU4g2CE318

# Walter Reed Doctor: First-Ever Powered Prosthesis:

http://www.youtube.com/watch?v=zZlmAvvYUi4

# Deputy Shot by Coworker Gets Bionic Ankle – WLNS:

http://www.youtube.com/watch?v=8z3-EhMLrCI

# My Experience with the IWalk BiOM:

http://www.youtube.com/watch?v=99k2hkb8Ylw

# iWalk BiOM Power Foot tested by Amputees in their 20's, 30's, 40's and 50's!:

http://www.youtube.com/watch?v=UfrnrCalbSY

# The $50,000 bionic leg:

http://www.youtube.com/watch?NR=1&v=uF--jvMZGUk&feature=endscreen

Bionic Soldiers: http://www.youtube.com/watch?v=52tFxg0O_bs

# Combat Vets Try Out Latest in Prosthetic Knee Design:

http://www.youtube.com/watch?v=N5rJCSFcBL8

# X2 Prosthetic Knee:

http://www.youtube.com/watch?v=kYXrd-smZb4

# Taylor Morris quadruple amputee, double above knee walks with Otto Bock:

http://www.youtube.com/watch?v=1JDyvlFEiPA

# Jeff Moore, Above Knee Amputee, Gives the X2/Genium Bionic Knee a Thumbs Up!:

http://www.youtube.com/watch?v=gL0W7T_wEB4

# X2 + BiOM foot = Astounding results!:

http://www.youtube.com/watch?v=BJx9pFa2Fec

# Landmine Victim Lebanon:

http://www.youtube.com/watch?v=X_EAXqk4MPQ

# The DEKA Arm:

http://www.youtube.com/watch?v=8FEkDDbmoVQ

# Dean Kamen's Robotic "Luke" Arm:

http://www.youtube.com/watch?v=R0_mLumx-6Y

# Claudia Mitchell Operates a Bionic Arm with her Brain at RIC:

http://www.youtube.com/watch?v=Aj7T-DZAdik

# Cutting Edge Prosthetic Arms:

http://www.youtube.com/watch?v=T6R5bm6qx2E

# Creating The Bionic Arm:

http://www.youtube.com/watch?v=6rloSSqiUCM

# Jesse Sullivan Operates a Bionic Arm with his Brain at RIC (no sound):

http://www.youtube.com/watch?v=ddInW6sm7JE&playnext=1&list=PLU7KpjB49o2_B658wZCG0Q0DSqkm9smgd&feature=results_main

# Amanda Kitts Operates a Bionic Arm with her Brain at RIC (no sound):

http://www.youtube.com/watch?v=qyCLuVOmZxo&list=PLU7KpjB49o2_B658wZCG0Q0DSqkm9smgd

# Kamal Walk with Jaipur Knee Aug 21 2008:

http://www.youtube.com/watch?v=bwnylcqA3g8